## Sheet2
SR Create Date – PE | PE - PLI # | Product – PE PLI | Lot Number – PE PLI | Mfg Site Name – PE PLI | Sub Business – PE PLI PM | Product Group – PE PLI PM | Product Description – PE PLI | Product Family Desc – PE PLI PM | Product Family Code – PE PLI PM | Product ID – PE PLI PM |  |  | NA | RFR | RFR Code Desc | AFC | AFC Code Desc | Summary of Investigation Results | Product Event ID | Legacy System ID – PE | Complaint? – PE | Country – PE | Status - PE | Event Description – PE | Facility - PE | Facility ID - PE | Notified Date – PE | Promotion Date – PE | Adv. Event Type – PE – Legacy | Event Date – PE | Event Context – PE | Closed date – PE | Outcome to AE - Congenital Anomaly – PE | Outcome to AE - Death – PE | Outcome to AE - Disability – PE | Outcome to AE - Hospitalization – PE | Outcome to AE - Intervention Required – PE | Outcome to AE - Life Threatening – PE | Outcome to AE - Other Description – PE | Outcome to AE - Other – PE | Event Summary – PE | RGA / CRGA – PE | Therapy – PE | Patient Outcome - PE | Interface Details – PE | Person Resp - PE | Analyst - Analysis PE | Created Date - Analysis PE | Completed Date - Analysis PE | Analysis Summary - Analysis PE | Status - Investigation | Reportable?
2018-03-01 | 0702399139-10 | 8881112599 | 17K18963X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:BURNED MATERIAL | Note creator: Daniel Callahan Note Date: 05/14/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. No physical sample was provided as of 05/14/18, but an image was received. The two (2) samples shown appear to have black discoloration of the barrel, with a smoky like texture. Without the physical sample, it can only be induced that the discoloration is embedded within the plastic material. A review of the device history record for Lot #17K18963X was performed with no discrepancies observed related to this defect. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112599 was 499,200 parts (800 parts per case), produced on Line 1 on 11/06/17 – 11/09/17. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The assessment eval uated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). Resin was burned inside the screw of the injection molding process, inherent part of manufacturing. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition is low risk and an inherent part of manufacturing. This complaint will be used for tracking and trending purposes. | 0702399139 | 0 | Y | United States | Closed | The customer reports: The syringes with what appears to be embedded burnt resin in the syringe. | NURSE ASSIST INC | 0012555852 | 2018-03-01 | 2018-03-14 | 0 | 2018-03-01 | During use | 2018-05-24 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: lot # 17K18963X Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event lot # 17K18963X syringes with what appears to be embedded burnt resin in the syringe Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: lot # 17K18963X As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-05 | 0702405459-10 | 8938 | 162271 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8938 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 162271 (10/13/14 to 10/20/14) and no internal reject found during production for the same issue. Trending Period: March 05, 2017 to March 23, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and sixteen other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8938 is 5 eaches/case. As per the complaint attachment 6 eaches are missing lid which show there was breakdown of product case in the supply chain. The case may be broken and repackaged outside of Cardinal Health for further distribution. This may have resulted in the failure of placing proper quantity of lids by prior to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702405459 | 0 | Y | United States | Closed | Customer received sharps container with missing lids. | 0 | 0 | 2018-03-05 | 2018-03-06 | 0 | 2018-03-05 | During use | 2018-03-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cardoso, Alexandra PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8938, Serial/Lot No:162271, Description:8938 RED 18 GAL CONT W SLIDING LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8938 RED 18 GAL CONT W SLIDING LID X5 Model Number: 8938 Serial or Lot Number: 162271 ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Customer received sharps container with missing lids. PHYSICIAN LAST NAME: Buck PHYSICIAN FIRST NAME: Diane PHYSICIAN CLINIC: Pathology Labratory PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 830 Bayou Pines West PHYSICIAN CITY: Lake Charles PHYSICIAN STATE CODE: TX PHYSICIAN POSTAL CODE: 70601 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: Pathology Labratory FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 830 Bayou Pines West FACILITY CITY: Lake Charles FACILITY STATE CODE: TX FACILITY POSTAL CODE: 70601 SALESREP USER ID: cardoa8 SALESREP EMP ID: 285356 SALESREP NAME: Cardoso, Alexandra PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cardoso, Alexandra As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-07 | 0702410855-10 | 89651 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 89651 2GAL RED CONT H DROP CLR LIDX20 | MITG - Sharps Disposal | 002927 | 89651 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 89651 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: March 7, 2017 to March 23, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the container/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 89651 is 20 eaches/case. As per complaint product line item the quantity is 3 eaches. The DHR cannot be reviewed as lot number is unknown. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702410855 | 0 | Y | United States | Closed | The customer reported Container arrived with missing lid | 0 | 0 | 2018-03-07 | 2018-03-08 | 0 | 2018-03-07 | During use | 2018-03-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cardoso, Alexandra PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:89651, Description:89651 2GAL RED CONT H DROP CLR LIDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 89651 2GAL RED CONT H DROP CLR LIDX20 Model Number: 89651 Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Container arrived with missing lid Was the customer notified the product should be returned? No Please explain n/a PHYSICIAN LAST NAME: Brown PHYSICIAN FIRST NAME: Leon PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Austin PHYSICIAN STATE CODE: TX For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: Community Clinical Research FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 8334 Cross Park Dr FACILITY CITY: Austin FACILITY STATE CODE: TX FACILITY POSTAL CODE: 78754 SALESREP USER ID: cardoa8 SALESREP EMP ID: 285356 SALESREP NAME: Cardoso, Alexandra PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cardoso, Alexandra As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-09 | 0702414496-10 | 8881512878 | 17H03463X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | PACKAGING | AFC | KC:DIFFICULT TO OPEN PACKAGE | Note creator: Daniel Callahan Note Date: 04/10/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 04/10/18, no samples or pictures have been provided. With no samples or pictures, we are unable to observe the reported condition. A review of the torque testing results and device history record for Lot #17H03463X was performed with no discrepancies observed. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. A defect could not be determined at this time. Torque testing is performed throughout the lot to ensure opening specifications are met. Within complaint, the customer confirmed being unfamiliar with sterile packaging. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition is low risk. In addition, customer was unfamiliar with new packaging. This complaint will be used for tracking and trending purposes. | 0702414496 | 0 | Y | Canada | Closed | The customer reports: Apparently Covidien change the presentation of the product ,now it is in a plastic box really difficult to open. No injury. | CARDINAL HEALTH CANADA INC RDP | 0012702795 | 2018-03-08 | 2018-03-09 | 0 | 2018-03-06 | Prior to Use | 2018-04-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-09 | 0702416435-10 | 8980 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8980 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: March 9, 2017 to March 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the container/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8980 is 10 eaches/case. The DHR cannot be reviewed as lot number is unknown. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702416435 | 0 | Y | United States | Closed | The customer reported 10 missing lids on item 8980. | SOUTHWEST GENERAL HEALTH CENTER | 0012919707 | 2018-03-09 | 2018-03-13 | 0 | 2018-03-09 | During use | 2018-03-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Unknown - DBS | 0 | FIRST NOTIFIED NAME: Platenyk, Kyle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for Neuro - DBS ------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8980, Description:8GAL CONTAINER NEEDLE DISP X10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8GAL CONTAINER NEEDLE DISP X10 Model Number: 8980 Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------- Regulations require that as much information as is available to Medtronic is reported. Help us to avoid follow-up communications by being thorough and clear in your descriptions of what happened. If information is unknown, please indicate why it is unknown or follow-up may be required. What is the issue being reported including the signs or symptoms the patient is experiencing related to the issue? (ex. Premature battery depletion, overdischarge, symptom return, withdrawal symptoms, pump alarm, increase in post implant pain) 10 missing lids on item 8980 Please describe any environmental/external/patient factors that may have led or contributed to the issue (ex. falls, EMI, patient compliance) no factors Please describe diagnostics/troubleshooting performed, include results and dates when known. If no diagnostics/troubleshooting performed, please state that below. no diagnostics Please describe interventions/actions that were taken to resolve the issue. If no interventions/actions were taken, please state that below. no actions taken Is the issue resolved at the time of this report? No Will the healthcare professional (HCP) have any further information on this event? No, the HCP has no further information for Medtronic SALESREP USER ID: platek1 SALESREP EMP ID: 326455 SALESREP NAME: Platenyk, Kyle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Platenyk, Kyle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-12 | 0702419040-10 | 8985 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985 8 GAL CHEMO CONT HINGED YEL X10 | MITG - Chemo Protection Safety | 003209 | 8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8985 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: March 12, 2017 to March 28, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the container/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8985 is 10 eaches/case. As per the complaint description the product was in the repack box. This shows that; there was breakdown of product case in the supply chain third party or distributor. As the information provided is very limited and there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702419040 | 0 | Y | United States | Closed | The customer reported they are missing the Hinged top, all we got was the container. No other item were included in this repack box | Integris Baptist Medical Ctr | 0008052690 | 2018-03-08 | 2018-03-12 | 0 | 0 | Prior to Use | 2018-04-04 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-13 | 0702421308-10 | 8985 | 17K12863 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985 8 GAL CHEMO CONT HINGED YEL X10 | MITG - Chemo Protection Safety | 003209 | 8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8985 evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 17K12863 (11/30/17 to 12/08/17). There was no internal reject found during production. Trending Period: March 12, 2017 to April 5, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing absorbent pads product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702421308 | 0 | Y | United States | Closed | The customer reported: The absorbant that is included in the case was missing . No damage to case. | CGH Medical Center | 0015922922 | 2018-03-12 | 2018-03-13 | 0 | 0 | Unknown | 2018-04-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-14 | 0702424615-10 | 8507SA | 17K15663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | LOW PUNCTURE FORCE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 8507SA evaluated for the contaminated needle stick. The reported condition was observed in the photo as needle was protruding in the middle of the container’s short side. DHR Review was performed and met the acceptance criteria. DHR Review: Device History Records data reviewed for 17K15663 (12/04/17 to 12/09/17) and no internal reject found during production. Trending Period: March 14, 2017 to April 06,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the same failure for the product family. Root Cause: The potential root cause associated with this event is improper handling/carrying of a filled sharps container below recommended fill line by user. As per complaint product analysis the container was punctured in the middle part of container’s short side. If sample will be available then further testing can be performed. It cannot be determined the fill level content from the photo provided. As per the IFU carrying/lifting instructions “Never hold or carry used container below recommended fill line. Use additional caution with containers that do not have handles.” Also per IFU precaution “do not hold or carry container close to the body”. Also there is a caution on the label stating “Forcing could cause puncture”. The excessive force may be used to drop sharps in the container resulted in the container puncture. The DHR was reviewed and no internal reject found. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The carrying/lifting instructions were not followed by user. Also the fill level of the container was not determined from the photo provided. | 0702424615 | 0 | Y | United States | Closed | The customer states that there is a needle went throught the plastic (sharps container) and injured the physician. The needle stick occurred on the physician's left hand. | 0 | 0 | 2018-03-14 | 2018-03-14 | 0 | 2018-03-13 | During use | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Serious Illness Injury | 0 | Daniel Tamul | Gatha Kothari | 43196 | 43196 | The photo was available for product analysis. The needle was protruding through container in the middle part of the container's short side. Unfortunately from the available photo it cannot be determined the fill level of the container. If sample will be available then further testing can be performed. | Complete | Yes
2018-03-15 | 0702426921-10 | 8881112059 | 17L02363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:INK | Note creator: Daniel Callahan Note Date: 04/20/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. Three (3) samples were received, and were outside original packaging per product specifications. Black particulate is located on each sample on the outside of the barrel. The locations of particulate are not identical for each syringe. A review of the device history record for Lot #17L02363X was performed with no discrepancies observed. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112059 was 144,000 parts (500 parts per case), produced on Line 23 on 11/15/17 – 11/17/17. Probable Root Cause: Doctor Blade and/or wiper blade were worn and not able to consistently prevent excess ink from building up on the cliché, which transferred to the barrel during printing. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition is low risk. See following Corrective Action: Updated Job Plan AM-256-2W-PM to include the following task on 12/07/17: Replace Doctor Blade, Wiper and Seals according to Flex Apex Procedure (bi-monthly) | 0702426921 | 0 | Y | United States | Closed | The customer reports: Production team reported syringes with black ink stains. No patient involved.-###-From PE-702435987-The customer reports: Ink smears on the outside of the barrel. No patient involved. | 0 | 0 | 2018-03-15 | 2018-03-16 | 0 | 2018-03-15 | Unknown | 2018-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas Event Context: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: CUSTOMER AT CRBARD - OMAR CALVARIO ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: 420 PIECES DURING 100% INSPECTION WERE FOUND DEFECTIVE. 7500 PIECES ARE UNDER SUSPECT. Below are the details for Product - Model No:8881112059, Serial/Lot No:17L02363X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: CUSTOMER - CRBARD CONTACT NAME AND PHONE: OMAR CALVARIO 1+956 205 7271 REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 17L02363X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event Production team reported syringes with black ink stains PICTURES ATTACHED Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Park, Lynn ADDITIONAL COMMENTS ABOUT THE EVENT: PLEASE SEND A DETAILED INVESTIGATION REPORT TO OMAR.CALVARIO@CRBARD.COM AND DOUGLAS.TOUSSAINT@ MEDTRONIC.COM AND LYNN.A.PARK@MEDTRONIC.COM As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes -###-From PE-702435987- FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Serial/Lot No:17L02363X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 17L02363X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event our customer found ink smears on the outside of the barrel. Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: ink smears on the outside of the barrel. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | Christopher Gildea | 43181 | 43181 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2018-03-16 | 0702429771-10 | 8881412012 | 17M07763X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881412012 CURVED TIP SYRINGE X50 | MITG - Needles & Syringes | 003208 | 8881412012 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:DEFORMED/WARPED | Note creator: Daniel Callahan Note Date: 05/01/18 Note Description: Addition of Investigation Summary As of 05/01/18, no samples or pictures of the product that caused the customer complaint have been provided. A review of the device history record was performed for Lot# 17M07763X, with no anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The assessment eval uated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). New Tool #58203 may have introduced special variation in the process that was not detected in addition to customer re-using product. 8881412012 is a 1x use product that cannot be re-sterilized. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition cannot be confirmed to be a manufacturing defect and the risk is minor. This complaint will be used for tracking and trending purposes. | 0702429771 | 0 | Y | Canada | Closed | The customer reports: The syringes were sterilized and it modified the tip making irrigation impossible. Did not happen to all syringes but to many of them from this lot. | 0 | 0 | 2018-03-16 | 2018-03-21 | 0 | 2018-03-16 | Prior to Use | 2018-05-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Lafortune, Dino PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Dental ---------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Sterilizatio tehcnician ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881412012, Serial/Lot No:17M07763X, Description:8881412012 CURVED TIP SYRINGE X50 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881412012 CURVED TIP SYRINGE X50 Model Number: 8881412012 Serial or Lot Number: 17M07763X ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event For this lot # syringes were sterilized and it modified the tip making irrigation impossible. Did not happen to all syringes but to many of them from this lot Was the customer notified the product should be returned? Yes FACILITY NAME: Fountainview Oral Surgery FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 4288 Center Court FACILITY CITY: Mississauga FACILITY STATE CODE: ON FACILITY POSTAL CODE: L4Z 1S2 FACILITY DAY PHONE: 905-275-0058 SALESREP USER ID: laford1 SALESREP EMP ID: 252160 SALESREP NAME: Lafortune, Dino PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Lafortune, Dino As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-16 | 0702430151-10 | 31378089 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31378089 19 GAL SHARPS W/GASKET REDX5 | MITG - Sharps Disposal | 002927 | 31378089 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31378089 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: March 15, 2017 to April 5, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaints reported for the product family container size/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 31378089 is 5 eaches/case. DHR cannot be reviewed as lot no. is not provided. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702430151 | 0 | Y | United States | Closed | The customer reported they received three cases of product with missing lids. | Elkhart Clinic | 0011803337 | 2018-03-15 | 2018-03-19 | 0 | 0 | During use | 2018-04-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:31378089, Description:31378089 19 GAL SHARPS W/GASKET REDX5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 31378089 19 GAL SHARPS W/GASKET REDX5 Model Number: 31378089 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Account received three cases of product with missing lids Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: pickrell PHYSICIAN FIRST NAME: lisa PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: elkart PHYSICIAN STATE CODE: IN For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-20 | 0702435987-10 | 8881112059 | 17L02363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | Duplicate | #N/A | 0702435987 | 0 | N | United States | Closed | The customer reports: Ink smears on the outside of the barrel. No patient involved. | BARD SHANNON LIMITED | 0012818702 | 2018-03-20 | 2018-03-21 | 0 | 2018-03-20 | During use | 2018-04-11 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Serial/Lot No:17L02363X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 17L02363X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event our customer found ink smears on the outside of the barrel. Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: ink smears on the outside of the barrel. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-23 | 0702444007-10 | 1181200777 | 17H02763X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1181200777 NZ ONLY 12ML LUER LOCK SYRING | MITG - Needles & Syringes | 003208 | 1181200777 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:SYRINGE LEAKING | Note creator: Daniel Callahan Note Date: 04/24/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 4/24/18, no samples or pictures of the product that caused the customer complaint have been provided. A review of the device history record was performed for Lot# 17H02763X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 1181200777 was 230,400 parts (480 parts per case), produced on Line 8 on 08/15/17-08/18/17 The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. Possible root cause was determined to be jams in the dials after printing caused a crack in the barrel that resulted in the customer complaint. Based on the information available, trend analysis, CAPA decision tool, and the investigation findings, a formal investigation is not deemed necessary at this time. The reported condition could not confirm a manufacturing defect due to absence of sample or image for analysis. | 0702444007 | 0 | Y | United States | Closed | The customer reported: A Covidien 12ml monoject syringe had a crack down the entire barrel of the syringe. Pharmacy technician while drawing up concentrated potassium chloride had solution spray all over her person. No injury occurred to technician. The medication was prepared again and dispensed to the patient. | Cape Fear Valley Medical Center | 0015924168 | 2018-03-23 | 2018-03-23 | 0 | 2018-01-01 | Unknown | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | Daniel Callahan | 43221 | 43236 | Note Creator: Daniel Callahan Note Date: 05/04/18 Note Description: Addition of analysis Sample received and analyzed on 05/04/18. Sample was received inside soft-pack packaging, with seal broken. Barrel contains a crack from the shoulder to the 9mL mark. No visible damage to the soft-pack packaging. Due to breakage of grad print on the crack, barrel was cracked after the pad printing station. Post-printing breakage correlates aligns with probable root cause. No change to investigation results. | Complete | Yes
2018-03-26 | 0702447026-10 | 8998PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998PG2 18 GAL RED PG2 HINGE LID X5 | MITG - Sharps Disposal | 002927 | 8998PG2 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8998PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: March 23, 2017 to April 17, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seventeen other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8998PG2 is 5 eaches/case. The lot number is unknown so DHR cannot be reviewed for manufacturing of the product. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702447026 | 0 | Y | United States | Closed | The customer reported missing one lid in case of sharps containers. | NORTHSIDE HOSP | 0016434865 | 2018-03-23 | 2018-03-28 | 0 | 2018-03-23 | During use | 2018-04-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8998PG2, Description:8998PG2 18 GAL RED PG2 HINGE LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8998PG2 18 GAL RED PG2 HINGE LID X5 Model Number: 8998PG2 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Missing one lid in case of sharps containers. PHYSICIAN LAST NAME: Coble PHYSICIAN FIRST NAME: Monica PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Cumming PHYSICIAN STATE CODE: GA PHYSICIAN POSTAL CODE: 30041 PHYSICIAN DAY PHONE: 770-205-1331 PHYSICIAN EMAIL: Monica.Coble@gacancer.com For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-03-29 | 0702456165-10 | 85131- | 17K22463 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85131 5QT RED CONT SHARPS M/BOX X20 | MITG - Sharps Disposal | 002927 | 85131- |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 85131 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17K22463 (01/11/18 to 01/17/18) and no internal reject found during production for the same issue. Trending Period: March 23, 2017 to April 17, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the container size/different lid style for 5 Qt. product family. There was no other complaint reported for 85131 with WP8510 lid. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 85131 is 20 eaches/case. There was no information in the complaints out of 20 lids how many were missing in the case. The DHR was reviewed and it passed all the internal inspections for the lot number. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702456165 | 0 | Y | United States | Closed | The customer reported they received case with missing lids. | Freeman Regional Health Services | 0004811869 | 2018-03-23 | 2018-03-30 | 0 | 2018-03-23 | Servicing | 2018-04-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Charlesworth, Victoria PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Supply Chain ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:85131-, Serial/Lot No:17K22463, Description:85131 5QT RED CONT SHARPS M/BOX X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 85131 5QT RED CONT SHARPS M/BOX X20 Model Number: 85131- Serial or Lot Number: 17K22463 ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event received case with missing lids SALESREP USER ID: charlv2 SALESREP EMP ID: 278791 SALESREP NAME: Charlesworth, Victoria PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-03 | 0702463444-10 | 8998PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998PG2 18 GAL RED PG2 HINGE LID X5 | MITG - Sharps Disposal | 002927 | 8998PG2 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8998PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 3, 2017 to April 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eighteen other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8998PG2 is 5 eaches/case. The lot number is unknown so DHR cannot be reviewed for manufacturing of the product. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702463444 | 0 | Y | United States | Closed | The customer reported missing lids when they received case. | GA CANCER SPECIALISTS ALPHARETTA | 0012969838 | 2018-04-03 | 2018-04-04 | 0 | 2018-04-03 | Prior to Use | 2018-05-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8998PG2, Description:8998PG2 18 GAL RED PG2 HINGE LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8998PG2 18 GAL RED PG2 HINGE LID X5 Model Number: 8998PG2 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Missing lids when customer received case PHYSICIAN LAST NAME: Ebert PHYSICIAN FIRST NAME: Sondra PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Alpharetta PHYSICIAN STATE CODE: GA PHYSICIAN DAY PHONE: 678-566-6995 PHYSICIAN EMAIL: Sondra.Ebert@gacancer.com For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-04 | 0702466062-10 | 8935 | 18C06063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8935 12GAL RED CONTAINERNEEDLEDISPX10 | MITG - Sharps Disposal | 002927 | 8935 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:CRACKED CONTAINER | The Product 8935 evaluated for the cracked container. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for lot number 18C06063 (03/24/18 to 03/28/18) and no internal reject found during production. Trending Period: April 4, 2017 to April 30, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The DHR was reviewed and it was packaged per specification and met acceptance criteria. The most probable root cause is the containers may have cracked during shipping and handling due external stress/forces on the container. The physical sample or photo were not available for product analysis. There was very minimal information in the complaint. It does not have any information about if the product was repackaged outside of Cardinal Health or not. As per complaint it was empty container and the issue was found before use. So there was no safety or harm to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. As per complaint it was empty container and the issue was found before use. So there was no safety or harm to end user. | 0702466062 | 0 | Y | United States | Closed | Customer reported that the container was cracked. | 0 | 0 | 2018-04-04 | 2018-04-04 | 0 | 2018-04-03 | Unknown | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2018-04-04 | 0702466131-10 | 8989PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8989PG2 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 4, 2017 to April 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eighteen other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989PG2 is 5 eaches/case. The lot number is unknown so DHR cannot be reviewed for manufacturing of the product. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702466131 | 0 | Y | United States | Closed | The customer reported Missing 5 lids in case. | GEORGIA CANCER SPCLST ATHENS | 0012730361 | 2018-04-04 | 2018-04-05 | 0 | 2018-04-04 | During use | 2018-05-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989PG2, Description:8989PG2 18 GAL CHEMO YEL PG2 HINGE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID Model Number: 8989PG2 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Missing 5 lids in case PHYSICIAN LAST NAME: Barnes PHYSICIAN FIRST NAME: Dee PHYSICIAN CLINIC: GCS Athens PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 125 King Ave PHYSICIAN CITY: Athens PHYSICIAN STATE CODE: GA PHYSICIAN POSTAL CODE: 30306 PHYSICIAN DAY PHONE: 706-369-4478 PHYSICIAN EMAIL: deena.barnes@gacancer.com For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-05 | 0702469158-10 | 8602RC | 17K38163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | (D) 2GA RCRA HAZARDOUSWASTE CONTX20 | MITG - Sharps Disposal | 002927 | 8602RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING ABSORBENT PADS | The Product 8602RC evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 17K38163 (02/02/18 to 02/04/18). There was a NCR 18020004 for missing kit component during QA Case inspection found during production. All affected product was contained and reworked as procedure and passed AQL sampling inspection. Trending Period: April 5, 2017 to April 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing absorbent pads product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702469158 | 0 | Y | United States | Closed | The customer states the absorbent pads were missing. | HENRY FORD WEST BLOOMFIELD | 0012604510 | 2018-04-05 | 2018-04-09 | 0 | 2018-04-05 | Servicing | 2018-05-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8602RC, Serial/Lot No:17K38163, Description:(D) 2GA RCRA HAZARDOUSWASTE CONTX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for (D) 2GA RCRA HAZARDOUSWASTE CONTX20 Model Number: 8602RC Serial or Lot Number: 17K38163 ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Received case with missing absorbent pads. SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-06 | 0702472591-10 | 8881112599 | 17M16263X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING COMPONENTS | Note creator: Daniel Callahan Note Date: 05/15/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 5/15/18, no samples or pictures of the product that caused the customer complaint have been provided. Without samples or pictures, we are unable to observe the reported condition and perform product analysis. A review of the device history record for Lot #17M16263X was performed with no discrepancies observed related to this defect. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112599 was 499,200 parts (800 parts per case), produced on Line 1 on 01/05/18 – 01/08/18. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes. The most probable root cause is last check sensors for missing plungers not functioning as intended. Based on the information available, and the investigation findings, a formal investigation is not deemed necessary at this time. Corrective action already identified for similar defect per Complaint #702330073. The following corrective actions are being implemented: 1. Updated Line 1 Checklist DC0006949 through Change Order # EC100737. Final Inspection sensors (plunger and tip presence) will be added to be inspected at beginning and end of each Lot. 2. Update PCR tied to Line 1 Checklist through Change Order #EC105021. | 0702472591 | 0 | Y | United States | Closed | The customer reports: missing plunger rods | NURSE ASSIST INC | 0012957619 | 2018-04-06 | 2018-04-09 | 0 | 2018-04-06 | During use | 2018-10-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Nurse Assist found 20 syringes missing plungfer rods. Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event missing plunger rods LOT 17M16263X Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: missing plunger rod LOT NUMBER 17M162623X As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-06 | 0702472606-10 | 8881112985 | 17M13963X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 10ML STERILE WATER SYR BULK NS | MITG - Needles & Syringes | 003208 | 8881112985 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:NO FAILURE | Note creator: Daniel Callahan Note Date: 05/15/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 5/15/18, no samples or pictures of the product that caused the customer complaint have been provided. Without samples, we are unable to investigate potential product issues that impact conductivity. A review of the device history record for Lot #17M13963X was performed with no discrepancies observed related to this defect. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112985 was 700,000 parts (1000 parts per case), produced on Line 7 on 01/05/18 – 01/11/18. No root cause can be determined. It is recommended customer reviews sterilization process for any special variation that may have occurred. Based on the information available, and the investigation findings, a formal investigation is not deemed necessary at this time as the reported condition is not a quality design requirement. It is recommended customer reviews sterilization process for special variation during Lot. In the event our customer would like to develop a custom product solution, it is recommended to contact a sales rep. | 0702472606 | 0 | Y | United States | Closed | The customer reports: Conductivity limits for the filled, sterile syringes are less than or equal to 25. Finished product testing for that lot ranged from 28-29. Filled syringes were tested prior to sterilization, and the results ranged from 0.36-0.65 as expected. Even though the in-process conductivity results met specification, they ended up significantly higher than usual, indicating something with the syringe changed or was different. | NURSE ASSIST INC | 0012957619 | 2018-04-06 | 2018-04-09 | 0 | 2018-04-06 | During use | 2018-10-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Conductivity limits for the filled, sterile syringes are less than or equal to 25. Finished product testing for that lot ranged from 28-29. Filled syringes were tested prior to sterilization, and the results ranged from 0.36-0.65 as expected. Even though the in-process conductivity results met specification, they ended up significantly higher than usual, indicating something with the syringe changed or was different. Below are the details for Product - Model No:8881112985, Description:10ML STERILE WATER SYR BULK NS IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 10ML STERILE WATER SYR BULK NS Model Number: 8881112985 Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------------- Please provide a detailed description of the event LOT NUMBER 17M13963X Conductivity limits for the filled, sterile syringes are less than or equal to 25. Finished product testing for that lot ranged from 28-29. Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: Conductivity limits for the filled, sterile syringes are less than or equal to 25. Finished product testing for that lot ranged from 28-29. Filled syringes were tested prior to sterilization, and the results ranged from 0.36-0.65 as expected. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-16 | 0702488686-10 | 5551500055 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | RUB TIP 5551500055 SMOOTH WEST 12ML | Unknown |  | 5551500055 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:DAMAGED COMPONENTS | Note creator: Daniel Callahan Note Date: 05/24/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was performed for Lot# 18B24563X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 5551500055 was 1,000,000 parts (10,000 parts per case), produced on 02/22/18. A root cause analysis was conducted in support of the investigation. The method of root cause analysis was to conduct a Six M assessment. The investigation confirmed that Covidien/Cardinal Health manufacturing processes could not contribute to the reported condition. Internal manufacturing processes use a purchased rubber tip which is introduced into rubber tip bowl that is designed with a smooth interior surface that is siliconized. Bowl inspection confirmed smooth surfaces are still intact. A siliconized smooth interior surface provides for little friction during further processing. Being a purchased component, the supplier was issued a SCAR #CAPA2018-014648 for investigation and implementation of corrective actions. The supplier has confirmed wear on one of their trim dies as the root cause of fringe. A product analysis was also conducted in support of the investigation. The customer reported fringe located on the rubber tip piston base pre-use as shown in the notification. Twenty-six (26) rubber tip samples were received in small plastic bag, and reported to have not been further processed at customer manufacturing facility. All samples were visually inspected and confirmed to have trim fringe, some without strings and some with a string closely attached to lower rib. Samples without strings were measured and found to be all within specification for max trim diameter. The customer reported that after assembly at their manufacturing site, the fringe was observed tearing off the rubber tip piston and creating string like extraneous matter as seen in the notification. String fringe is in the non-fluid pathway and behind the ribs of the rubber tip piston as shown in the notification. Rubber tip ribs are designed to a larger diameter than the max trim specification. It has been noted that the syringe is not a Covidien/Cardinal Health designed and supplied syringe. The syringe manufacturing process and syringe design can exacerbate rubber tip piston fringe. Internally produced rubber tip pistons and syringes are pre-siliconized to aid in assembly and ensure smooth actuation. Internally produced Syringes include inner diameters that account for max rubber tip piston trim diameter and include a taper from tighter (luer) to looser (finger flange). Customer manufacturing process and syringe design syringe shall be considered for evaluation. Twenty-one (21) of the supplied samples were assembled with internally produced and designed syringes and fully actuated. String fringe was not created using internally designed /manufactured syringes. Of the twenty-one (21) samples, only one (1) sample with a pre-existing string was observed during inspection. The string did not enter the syringe fluid pathway. ISO 80369-20:2016 positive pressure leak test was completed. No leakage was observed. Testing confirmed the reported condition has no patient or performance impact when using internal manufacturing processes and designed syringes. | 0702488686 | 0 | Y | Denmark | Closed | The customer reports: Plunger tips with threads - this quality issue was identified during the manufacturing process at Carmo plant. The devices were not sent to end users. | Carmo A/S - Hofstätter & Ebbesen A/ | 0016808718 | 2018-04-09 | 2018-04-16 | 0 | 0 | Prior to Use | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-18 | 0702495656-10 | 8989 | 17F05063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 89389 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17F05063 (08/12/17 to 08/23/17) and no internal reject found during production for the same issue. Trending Period: April 18, 2017 to May 11, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702495656 | 0 | Y | United States | Closed | Customer reports: there were no lids inside the box. | ATLANTA CANCER CARE | 0012906427 | 2018-04-18 | 2018-04-19 | 0 | 2018-04-18 | During use | 2018-05-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Chemo Protection Safety --------------------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989, Serial/Lot No:17F05063, Description:8989 18 GAL CONT CHEMO HNGD LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8989 18 GAL CONT CHEMO HNGD LID X5 Model Number: 8989 Serial or Lot Number: 17F05063 ------------------------------------------------------------------------ ------------------------------- When was the issue noticed? Before use/during setup Please provide a detailed description of the event The unopened case arrived with no lids PHYSICIAN LAST NAME: Farlow PHYSICIAN FIRST NAME: Kelly PHYSICIAN CLINIC: Atlanta Cancer Care PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Jonesboro PHYSICIAN STATE CODE: GA PHYSICIAN DAY PHONE: 404-851-2395 PHYSICIAN EMAIL: kfarlow@atlantacancercare.com For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-20 | 0702501496-10 | 8989PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8989PG2 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8989PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 20, 2017 to May 11, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989PG2 is 5 eaches/case. The packaging process is manual operation. The DHR cannot be reviewed as lot number is not provided. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702501496 | 0 | Y | United States | Closed | Customer reports: the lids are missing. | CANDLER HOSPITAL | 0013263368 | 2018-04-20 | 2018-04-23 | 0 | 2018-04-20 | During use | 2018-05-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989PG2, Description:8989PG2 18 GAL CHEMO YEL PG2 HINGE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID Model Number: 8989PG2 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The unopened case arrived with no lids For Physician with Siebel ID 0014662874: PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0014662874 Julie Foster cannot find a BP in GCH | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-23 | 0702503860-10 | 8507SA | 18B11163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING CONTAINER | The Product 8507SA evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B11163 (04/04/18 to 04/11/18) and no internal reject found during production for the same issue. Trending Period: April 20, 2017 to May 15, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the container size/lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify and follow packaging standard work and weight scale check for the 8507SA. The case quantity for 8507SA is 20 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing containers product should not be used. There was no patient involvement so no safety or harm to patient. | 0702503860 | 0 | Y | United States | Closed | The customer reported: When they opened the case they were missing 10 sharp containers. They had all 20 lids. Per custpmer there was not any damage or appear to be repacked. Customer looking for the 10 missing sharp containers. | OSF HEALTHCARE SYSTEMS-OWNED | 0012613179 | 2018-04-20 | 2018-04-23 | 0 | 0 | Unknown | 2018-05-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-23 | 0702503943-10 | 8991 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8991 RED 18GAL SHARPS DISP X5 | MITG - Sharps Disposal | 002927 | 8991 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8991 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 23, 2017 to Sep 10, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty nine other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8991 is 5 eaches/case. There is no information available in the complaint whether customer received full case or not. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702503943 | 0 | Y | United States | Closed | The customer reported: The container was missing 1 lid. | MEDLINE | 0012894606 | 2018-04-23 | 2018-04-23 | 0 | 0 | Unknown | 2018-09-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-24 | 0702506616-10 | 8932 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8932 RED 12 GAL CONT HINGE/SEAL LIDX10 | MITG - Sharps Disposal | 002927 | 8932 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8932 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 24, 2017 to May 18, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8932 is 10 eaches/case. As per the complaint attachment 2 case were missing 1 lid in each case. As there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702506616 | 0 | Y | United States | Closed | The customer reported: We have order 2 cases of KDL8932 two cases only had 9 gasket covers for the 10 container please replace 2 gasket covers if cannot replace please issue credit 2 each | Cheshire Medical Center | 0011804982 | 2018-04-24 | 2018-04-24 | 0 | 2018-04-05 | Prior to Use | 2018-05-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-30 | 0702517800-10 | 8980 | 910018B09063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:COSMETIC ISSUE | Reference attached investigation in GCH for a full concise summary. | 0702517800 | 0 | Y | United States | Closed | The Customer reported: Lids have burn marks.-###-From PE-702570014-The customer reports: Lids have burn marks. | HENRY SCHEIN | 0013106556 | 2018-04-27 | 2018-04-30 | 0 | 0 | Unknown | 2018-06-09 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-04-30 | 0702518463-10 | 8980 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8980 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 30, 2017 to Sep 10, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8980 is 10 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702518463 | 0 | Y | United States | Closed | The Customer reported: Missing lids | FISHER SCIENTIFIC | 0013266740 | 2018-04-30 | 2018-04-30 | 0 | 0 | Unknown | 2018-09-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-01 | 0702520321-10 | 8980- | 18B09063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980- |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | Duplicate | #N/A | 0702520321 | 0 | Y | United States | Closed | The Customer reported: Burn marks on lids | HENRY SCHEIN | 0012869883 | 2018-04-30 | 2018-05-01 | 0 | 0 | Unknown | 2018-05-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-03 | 0702526365-10 | 1181200777 | 17H02863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1181200777 NZ ONLY 12ML LUER LOCK SYRING | MITG - Needles & Syringes | 003208 | 1181200777 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:BROKEN LUER TIP | Note creator: Daniel Callahan Note Date: 07/02/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 07/02/18, no defect was reported and no sample was received. Investigation will be re-opened when sample is received and defect is observed. | 0702526365 | 0 | Y | United States | Closed | The customer reported issue: None; No further information provided. The sample was returned for evaluation. Based on the sample evaluation results, the luer tip is broken as seen in attached image. Magnification was used to view the luer surface around the broken tip for roughness, and was confirmed when gliding finger along the edge. Unknown liquid is present inside the syringe, which indicates attempted use at the customer location. | Cape Fear Valley Medical Center | 0015924168 | 2018-04-09 | 2018-05-03 | 0 | 0 | Unknown | 2018-08-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | Daniel Callahan | 43301 | 43301 | Note Creator: Daniel Callahan Note Date: 07/20/18 Note Description: Addition of analysis One (1) sample was received on 07/11/18. Sample has been removed from original packaging and soft-pack seal has been broken. The luer tip is broken as seen in attached image. Magnification was used to view the luer surface around the broken tip for roughness, and was confirmed when gliding finger along the edge. Unknown liquid is present inside the syringe, which indicates attempted use at the customer location. Customer provides no details of defect or the circumstances to when the defect was observed. No information was provided. | Complete | Yes
2018-05-03 | 0702526504-10 | 8870- |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGCOLLEC | 18GAL PHARMSAFETY WASTE CONT 5 | MITG - Waste Containers & Bags | 002897 | 8870- |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8870 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17K30263 (01/16/18 to 01/18/18) and no internal reject found during production for the same issue. Trending Period: May 3, 2017 to May 24, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty two other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8870 is 5 eaches/case. The packaging process is manual operation. As per the complaint the source is distribution. There may be potential risk of non-Cardinal Health distribution center breakdown the Cardinal Case and pack out in low unit of measure quantity for distribution to end user resulted in missing lids. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702526504 | 0 | Y | United States | Closed | The customer reports that 3 container lids were missing. | HCA Far West Las Vegas Distribution | 0001142444 | 2018-05-03 | 2018-05-10 | 0 | 2018-05-03 | During use | 2018-05-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | No Product found for LotSerial# 17K30263 FIRST NOTIFIED NAME: Daniel, Alan-Anthony PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distribution ASSOCIATED WITH PATIENT: No Below are the details for Product - Serial/Lot No:17K30263, Description:18 Gallon Gasket Container, Hinged Lid IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 18 Gallon Gasket Container, Hinged Lid Model Number: Not specified Serial or Lot Number: 17K30263 ------------------------------------------------------------------------ -------------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event A case of item 8870 was delivered with 3 lids missing FACILITY NAME: HCA Far West DC FACILITY COUNTRY CODE: USA FACILITY CITY: Las Vegas FACILITY STATE CODE: NV SALESREP USER ID: daniea17 SALESREP EMP ID: 349087 SALESREP NAME: Daniel, Alan-Anthony PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daniel, Alan-Anthony As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-07 | 0702530480-10 | 8507SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:COSMETIC ISSUE | The Product 8507SA evaluated for the cosmetic issue (grease). The reported condition was observed on the physical sample having grease on lid component of the lid assembly. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: May 7, 2017 to May 31, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error to inspect product for contamination and clean the part prior to packing for the finished good case. The root cause location might be injection molding machine, mold or conveyor had contaminant and it adhered to part during production. The DHR cannot be reviewed as no lot number information available. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as affected product should not be used. The product should be inspected prior to use for any abnormality. | 0702530480 | 0 | Y | Japan | Closed | The customer states, prior to use, a smear was found on the lid. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2018-05-07 | 2018-05-08 | 0 | 2018-05-06 | Prior to Use | 2018-07-26 | N | N | N | N | N | N | 0 | N | customer reports: there is a smear on the lid. | 0 | Null | 0 | 0 | Rose Matthews | Gatha Kothari | 43251 | 43251 | The physical sample was submitted with this complaint. The reported condition stated by the customer was observed. The grease contamination was on the sample between lid and hood area of lid assembly. The grease was adhered on the lid component. | Complete | #N/A
2018-05-11 | 0702542081-10 | 8506SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 5QT.CNT.W/COUNTR BLNC DR CLR20 | MITG - Sharps Disposal | 002927 | 8506SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8506SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: May 10, 2017 to June 4,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8506SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint description the quantity missing is 10 eaches. As per the attached report it was determined the product was from distributor. The case may be repackaged outside of Cardinal Health for further distribution to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702542081 | 0 | Y | United States | Closed | The Customer reported: That when they opened the CS there was only 10 EA in the CS. This should have 20 Ea per CS. | Rice Memorial Hospital | 0015739576 | 2018-05-11 | 2018-05-11 | 0 | 2018-05-10 | Prior to Use | 2018-06-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-22 | 0702562649-10 | 8507SA | 18A04163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18A04163 (02/19/18 to 02/27/18) and no internal reject found during production for the same issue. Trending Period: May 21, 2017 to June 12,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint information the quantity missing is 100 eaches and source is distributor. This seem like external distribution broken down and repack product to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702562649 | 0 | Y | United States | Closed | Customer reports: Lids were missing from the cases. | 0 | 0 | 2018-05-21 | 2018-05-23 | 0 | 2018-05-21 | During use | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: There were 272 lids missing amongst the three lot numbers, customer did not have an exact as to how many from each lot, as they combined everything from their shipment. Below are the details for Product - Model No:8507SA, Serial/Lot No:18A04163, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18A04163 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:17L23963, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 17L23963 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:18C01863, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18C01863 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-22 | 0702562649-20 | 8507SA | 17L23963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17L23963 (01/18/18 to 01/25/18) and no internal reject found during production for the same issue. Trending Period: May 21, 2017 to June 12,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint information the quantity missing is 100 eaches and source is distributor. This seem like external distribution broken down and repack product to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702562649 | 0 | Y | United States | Closed | Customer reports: Lids were missing from the cases. | 0 | 0 | 2018-05-21 | 2018-05-23 | 0 | 2018-05-21 | During use | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: There were 272 lids missing amongst the three lot numbers, customer did not have an exact as to how many from each lot, as they combined everything from their shipment. Below are the details for Product - Model No:8507SA, Serial/Lot No:18A04163, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18A04163 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:17L23963, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 17L23963 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:18C01863, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18C01863 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-22 | 0702562649-30 | 8507SA | 18C01863 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18C01863 (04/14/18 to 04/19/18) and no internal reject found during production for the same issue. Trending Period: May 21, 2017 to June 12,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint information the quantity missing is 72 eaches and source is distributor. This seem like external distribution broken down and repack product to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702562649 | 0 | Y | United States | Closed | Customer reports: Lids were missing from the cases. | 0 | 0 | 2018-05-21 | 2018-05-23 | 0 | 2018-05-21 | During use | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: There were 272 lids missing amongst the three lot numbers, customer did not have an exact as to how many from each lot, as they combined everything from their shipment. Below are the details for Product - Model No:8507SA, Serial/Lot No:18A04163, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18A04163 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:17L23963, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 17L23963 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases Below are the details for Product - Model No:8507SA, Serial/Lot No:18C01863, Description:8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 Model Number: 8507SA Serial or Lot Number: 18C01863 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the cases SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-24 | 0702568549-10 | V8536 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | VHA PLUS SSII 3GA CLR INRM 10 | MITG - Sharps Disposal | 002927 | V8536 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product V8536 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: May 24, 2017 to June 20,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for V8536 is 10 eaches/case. The manufacturing has manual packaging process. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702568549 | 0 | Y | United States | Closed | Customer reports: There were 2 missing container lids. | St. Tammany Hospital | 0004050517 | 2018-05-24 | 2018-05-29 | 0 | 2018-05-24 | During use | 2018-06-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cardoso, Alexandra PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:V8536, Description:VHA PLUS SSII 3GA CLR INRM 10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for VHA PLUS SSII 3GA CLR INRM 10 Model Number: V8536 Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Case of containers was received with 2 missing lids. PHYSICIAN LAST NAME: McKee PHYSICIAN FIRST NAME: Jodi PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Covington PHYSICIAN STATE CODE: LA For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP NAME: Isabell, Ashleigh PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cardoso, Alexandra As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-24 | 0702568937-10 | 8998PG2 | 18B20063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998PG2 18 GAL RED PG2 HINGE LID X5 | MITG - Sharps Disposal | 002927 | 8998PG2 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8998PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B20063 (03/27/18 to 03/29/18) and no internal reject found during production for the same issue. Trending Period: May 24, 2017 to June 20, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty two other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8998PG2 is 5 eaches/case. The manufacturing has manual packaging inspection. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702568937 | 0 | Y | United States | Closed | Customer reports: Lids were missing upon opening the package. | MCLAREN PORT HURON | 0013117647 | 2018-05-24 | 2018-05-29 | 0 | 2018-05-24 | Prior to Use | 2018-06-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8998PG2, Serial/Lot No:18B20063, Description:8998PG2 18 GAL RED PG2 HINGE LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8998PG2 18 GAL RED PG2 HINGE LID X5 Model Number: 8998PG2 Serial or Lot Number: 18B20063 ------------------------------------------------------------------------ ----------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing upon opening the package Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Denise PHYSICIAN FIRST NAME: Kivel PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Port Huron PHYSICIAN STATE CODE: MI For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-29 | 0702575490-10 | 8985S | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985S 8 GAL YEL CHEMO W/SLIDE LID X10 | MITG - Sharps Disposal | 002927 | 8985S |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8985S evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: May 29, 2017 to June 25,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to end user. The case quantity for 8985S is 10 eaches/case. As per the complaint the product was distributed by third party. This seem like external distribution broken down and repack product to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702575490 | 0 | Y | United States | Closed | Customer reported a missing lid. | St Marys Hospital | 0005787299 | 2018-05-29 | 2018-05-30 | 0 | 2018-05-29 | Servicing | 2018-06-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cohen, Brian PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Purchasing ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8985S, Description:8985S 8 GAL YEL CHEMO W/SLIDE LID X10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8985S 8 GAL YEL CHEMO W/SLIDE LID X10 Model Number: 8985S Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing SALESREP USER ID: cohenb4 SALESREP EMP ID: 328562 SALESREP NAME: Cohen, Brian PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cohen, Brian As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-29 | 0702575555-10 | 31353611 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 31353611 5 QRTG/GUARDIN-RMTRANSYEX14 | MITG - Sharps Disposal | 002927 | 31353611 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING CONTAINER | The Product 31353611 evaluated for the missing lids. The reported condition cannot be observed on photo provided. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18A22463 (04/23/18 to 04/26/18) and no internal reject found during production for the same issue. Trending Period: May 18, 2017 to June 8,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of containers prior to seal the case. The case quantity for 31353611 is 14 eaches/case. As packaging is manual operation. The DHR was reviewed and met acceptance criteria as required. The complaint description stated 15 lids and no container which does not match with case quantity of 14. There is not much information available for further investigation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing containers product should not be used. | 0702575555 | 0 | Y | Canada | Closed | Customer reports: The case contained only 15 lids with no containers in the box. | THOMPSON GENERAL HOSPITAL | 0012639554 | 2018-05-18 | 2018-05-30 | 0 | 2018-05-18 | Unknown | 2018-07-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cadoreth, Andre Event Context: Asked but unknown PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: email ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:31353611, Description:31353611 5 QRTG/GUARDIN-RMTRANSYEX14 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 31353611 5 QRTG/GUARDIN-RMTRANSYEX14 Model Number: 31353611 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event wallmounted containers came in and when we gave a case out they opened the case up only to find 15 lids and no containers in bottom of box Was the customer notified the product should be returned? Yes FACILITY NAME: Thompson Health Region FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 871 Thompson Drive South FACILITY CITY: Thompson FACILITY STATE CODE: MB FACILITY POSTAL CODE: R8N 0C8 FACILITY DAY PHONE: 204-6775318 FACILITY EMAIL: jhewitson@nrha.ca As Reported Event - departmentName: shipping SALESREP USER ID: russej2 SALESREP EMP ID: 331012 SALESREP NAME: Russell-Mondor, Joanne PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Russell-Mondor, Joanne ADDITIONAL COMMENTS ABOUT THE EVENT: customer care received email and sent it to the rep As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | Gatha Kothari | 43259 | 43259 | The photos of the 31353611 was attached in GCH. As from the photo lot number is 18A22463. The lids were present in the case. As from observed photos there were no containers in the case. | Complete | #N/A
2018-05-29 | 0702575998-10 | 8881112083 | 16M04263X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112083 SYR 112 S-PLK SP X800 | MITG - Needles & Syringes | 003208 | 8881112083 |  |  |  | RFR | NO INFORMATION | AFC | KC:MISSING PRINT | Note creator: Daniel Callahan Note Date: 07/02/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. Six (6) samples were received and visually inspected on 06/28/18. Samples do not appear to have been further processed. A review of the device history record was performed for Lot# 16M04263X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112083 was 144,000 parts (800 parts per case), produced on 01/29/17 – 02/01/17. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. Probable Root Cause: Rail was out of alignment causing gap between rail and transfer sprocket. This caused misaligned barrels under pad. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702575998 | 0 | Y | United States | Closed | The customer reported: Graduation marks are either missing or illegible. | QOSINA CORPORATION | 0012880103 | 2018-05-25 | 2018-05-29 | 0 | 2018-05-25 | Prior to Use | 2018-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Sunitha Krishnanpillai | Daniel Callahan | 43283 | 43283 | Note Creator: Daniel Callahan Note Date: 07/02/18 Note Description: Addition of analysis Six (6) samples were received and visually inspected on 06/28/18. Samples do not appear to have been further processed. Five (5) samples are missing one or more numerals, of these, one also has smeared graduation lines. Per GS10000795, these fall under Major B / 0.25 AQL. One (1) sample has illegible print for one numeral, which is a Major C / 0.40 AQL. Most samples contained illegible/missing numerals within the first 3mL on the barrel. | Complete | #N/A
2018-05-30 | 0702577935-10 | 31157160 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | FISHER 10 GA SHARPS CONT | MITG - Sharps Disposal | 002927 | 31157160 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 31157160 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: May 29, 2017 to June 25,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the product case. The case quantity for 31157160 is 6 eaches/case. The manufacturing has manual packaging. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702577935 | 0 | Y | United States | Closed | The customer reported: Received sharps containers with out lids. | Thermo Fisher Scientific | 0015890421 | 2018-05-29 | 2018-05-30 | 0 | 0 | Unknown | 2018-06-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-05-30 | 0702579078-10 | 8990Y | 00176813 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8990Y YEL 2 GAL BIOHAZ W/O GASKET X20 | MITG - Sharps Disposal | 002927 | 8990Y |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:DAMAGED PRODUCT | The Product 8990Y evaluated for the damaged product after autoclave process. The reported condition was observed in the photo provided. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176813 (5/13/17 to 5/15/17) and no internal reject found during production for the same issue. Trending Period: May 30, 2017 to June 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause may be higher temperature/duration for autoclave process used by customer. As per LS10000083 “Autoclave the container at 275ºF (135ºC) for one hour.” There was no information in the complaint about temperature and duration for the autoclave process used by customer. However, the use of an autoclave to decontaminate the contents of a Sharps Container is only appropriate when it is part of a complete system specifically designed for waste treatment. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event due to customer misuse. The autoclave temperature and duration used for disposal treatment used by customer is unknown. However, the use of an autoclave to decontaminate the contents of a Sharps Container is only appropriate when it is part of a complete system specifically designed for waste treatment. | 0702579078 | 0 | Y | Canada | Closed | Customer reports: There were two situations during the autoclave process, the sharps container and the yellow hinged lid on the top of the container melted. | PROVINCIAL TOXICOLOGY CENTRE | 0012911152 | 2018-05-30 | 2018-05-31 | 0 | 0 | Servicing | 2018-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Prokopy, Kimberly EVENT DATE: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8990Y, Serial/Lot No:00176813, Description:8990Y YEL 2 GAL BIOHAZ W/O GASKET X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8990Y YEL 2 GAL BIOHAZ W/O GASKET X20 Model Number: 8990Y Serial or Lot Number: 00176813 ------------------------------------------------------------------------ ----------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event This lab has encountered two situations where they went to autoclave the sharps container for disposal and the yellow hinged lid on the top of the container melted. The lot was the same for both containers. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Ng PHYSICIAN FIRST NAME: Monica PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Vancouver PHYSICIAN STATE CODE: BC PHYSICIAN POSTAL CODE: V5Z 4R4 PHYSICIAN DAY PHONE: 604-707-2672 PHYSICIAN EMAIL: monica.ng@bccdc.ca For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes FACILITY NAME: BCCDC Public Health Microbiology & Reference Lab FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 655 West 12th Ave FACILITY CITY: Vancouver FACILITY STATE CODE: BC FACILITY POSTAL CODE: V5Z 4R4 As Reported Event - departmentName: TB/Mycobateriology Laboratory SALESREP USER ID: prokok1 SALESREP EMP ID: 252524 SALESREP NAME: Prokopy, Kimberly PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Prokopy, Kimberly As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | Gatha Kothari | 43278 | 43278 | The photos were attached in the complaint. The yellow cap was melted on the lid assembly after autoclave process. | Complete | #N/A
2018-06-01 | 0702583738-10 | 85221R | 17L19863 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85221R 3GAL TRANS RED CONTAINER CLRTOP | MITG - Sharps Disposal | 002927 | 85221R |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 85221R evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17L19863 (02/22/18 to 03/10/18) and no internal reject found during production for the same issue. Trending Period: May 31, 2017 to June 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 85221R is 10 eaches/case. The manufacturing has manual packaging inspection. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient as this was discovered prior to use per complaint. | 0702583738 | 0 | Y | United States | Closed | The customer reported: Issue per customer: and that 2 containers were missing 2 lids | Cardinal Health | 0015754524 | 2018-05-31 | 2018-06-01 | 0 | 0 | Unknown | 2018-06-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-01 | 0702584042-10 | 8507SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING CONTAINER | The Product 8507SA evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: May 31, 2017 to June 26, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the container size/lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify and follow packaging standard work and weight scale check for the 8507SA. The case quantity for 8507SA is 20 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing containers product should not be used. There was no patient involvement so no safety or harm to patient. | 0702584042 | 0 | Y | United States | Closed | The customer reported: Received 20 lids, however they only received 10 containers | GIBSON COMMUNITY HOSP ASSOC | 0013132759 | 2018-05-31 | 2018-06-01 | 0 | 0 | Unknown | 2018-06-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-05 | 0702588955-10 | 8934CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8934CRS | =VLOOKUP(K48,'Cross Reference'!A:B,2,0) | =VLOOKUP(K48,'Cross Reference'!A:B,2,0) |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:LID FIT ISSUE | The Product 8934CRS evaluated for the lid fit issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: June 1, 2017 to June 28, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The lot number is unknown so Device History Record cannot be reviewed for 8934CRS. The most probable root cause per CAPA2016-010397 the lid R7417B was out of specification in lifter area and/or per CAPA2017-012361 is the R7413B container tool has a larger rim height than the R7413A tool, which causes the lid locks to be difficult to secure around the rim. Based on this CAPA all investigation is completed and corrective action and preventative actions are implemented and CAPA is closed. As there is no lot number provided it cannot be determine when the product was manufactured. Formal Investigation: Refer to CAPA 2016-010397 and CAPA2017-012361 for more details for corrective and preventive action. Conclusion: Refer to CAPA 2016-010397 and CAPA2017-012361 for more details for corrective and preventive action. | 0702588955 | 0 | Y | Canada | Closed | Customer reports: the lid is too tight to put on the container. | Cross Cancer Institute | 0001103819 | 2018-06-01 | 2018-06-19 | 0 | 0 | During use | 2018-11-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-06 | 0702592383-10 | 8602RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | (D) 2GA RCRA HAZARDOUSWASTE CONTX20 | MITG - Sharps Disposal | 002927 | 8602RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING COMPONENTS | The Product 8602RC evaluated for the missing components- waste labels. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record cannot be reviewed as lot number information is not provided. Trending Period: June 19, 2017 to June 28, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family for missing components in pre-kit label assembly. Root Cause: A probable root cause may be human error and failure to verify the amount of the “Hazardous Waste Labels” ” in the pre-kit assembly prior to sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing components – waste labels should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702592383 | 0 | Y | United States | Closed | Customer reported: Missing tampering labels on sharps containers | PATTERSON LOGISTICS SERVICES | 0012972160 | 2018-06-06 | 2018-06-06 | 0 | 2018-06-04 | Unknown | 2018-06-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-07 | 0702594339-10 | 8934CRS | 18B06063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8934CRS |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:LID FIT ISSUE | The Product 8934CRS evaluated for the lid fit issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B06063 (02/22/18 to 02/26/18) and no internal reject found during production for the same issue. Trending Period: June 1, 2017 to June 28, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The same customer has another complaint reported for 8934CRS PE 702588955 with the unknown lot number. The lot number for this complaint is 18B06063. The product was manufactured during 02/22/18 to 02/26/18. The most probable root cause may be customer had lids mixed up from unknown lot with good lot 18B06063. As there were previous CAPA for the lid fit issue addressed the root case. The most probable root cause per CAPA2016-010397 the lid R7417B was out of specification in lifter area and/or per CAPA2017-012361 is the R7413B container tool has a larger rim height than the R7413A tool, which causes the lid locks to be difficult to secure around the rim. Based on this CAPA all investigation is completed and corrective action and preventative actions are implemented and CAPA is closed. As lot number 18B06063 was manufactured in Feb 2018 after all CAPA actions are implemented so this seems like two different lot number were mixed at customer location. Formal Investigation: Refer to CAPA2016-010397 and CAPA2017-012361 for more details for corrective and preventive action. Conclusion: Refer to CAPA2016-010397 and CAPA2017-012361 for more details for corrective and preventive action. | 0702594339 | 0 | Y | Canada | Closed | Customer reports: the lids are extremely difficult to clip on to the container. | Cross Cancer Institute | 0001103819 | 2018-06-06 | 2018-06-19 | 0 | 2018-05-28 | Prior to Use | 2018-11-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-11 | 0702600461-10 | 8989 | 18B18763 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B18763 (03/29/18 to 04/02/18) and no internal reject found during production for the same issue. Trending Period: June 8, 2017 to July 2, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty two other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702600461 | 0 | Y | United States | Closed | The customer reported: 1 case was missing the lids for the containers | Lawrence Memorial Hospital | 0015738678 | 2018-06-08 | 2018-06-11 | 0 | 0 | Unknown | 2018-07-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-11 | 0702601340-10 | 8612RC | 00166570X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GA RCRA HAZARDOUS WASTE | MITG - Sharps Disposal | 002927 | 8612RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING COMPONENTS | The Product 8612RC evaluated for the missing components- kit item (absorbent pads/waste-toxic labels). The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 166570 (08/20/15-08/21/15). There was no internal reject found during production. Trending Period: June 11, 2017 to July 09, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the kit item (absorbent pads/waste-toxic labels) packed prior to sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing kit item (absorbent pads/waste-toxic labels) for product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702601340 | 0 | Y | United States | Closed | Customer reports: the case was missing the labels and pads. | SHARPS COMPLIANCE INC | 0012592746 | 2018-06-11 | 2018-06-14 | 0 | 2018-06-11 | During use | 2018-11-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Discovered By Distributor: Sharps Compliance Contact: Esmie Gonzalez 713-660-3534 egonzales@sharpsinc.com Below are the details for Product - Model No:8612RC, Description:12 GA RCRA HAZARDOUS WASTE IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 12 GA RCRA HAZARDOUS WASTE Model Number: 8612RC Serial or Lot Number: null ------------------------------------------------------------------------ --------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lot# (10) 00166570X, entire case was missing the labels and pads SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | Juan Diaz | 43299 | 43299 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2018-06-13 | 0702604950-10 | 8980- | 18B08663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980- |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:COSMETIC ISSUE | The Product 8980 evaluated for the cosmetic issue (burn mark). The reported condition was observed on the physical sample having burn mark on lid. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B08663 (02/15/18 to 03/05/18) and no internal reject found during production. Trending Period: June 13, 2017 to July 10, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The burn mark is Minor A class per GS10000913 cosmetic defect which does not affect form, fit and function for product intended use. As per inspection standard it has only reaction plan to minimize. The DHR was reviewed and met acceptance criteria. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as burn mark is the cosmetic issue which does not affect form, fit and function of product intended use. | 0702604950 | 0 | Y | Japan | Closed | Customer reports: there were black stains on the lids. | Niti-on Co., Ltd | 0015877231 | 2018-06-13 | 2018-06-13 | 0 | 2018-06-12 | Prior to Use | 2018-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Gatha Kothari | 43291 | 43291 | The physical sample was submitted with this complaint. The reported condition stated by the customer was observed. The burn mark is Minor A class per GS10000913 cosmetic defect which does not affect form, fit and function for product intended use. As per acceptance criteria it is acceptable. | Complete | #N/A
2018-06-18 | 0702613906-10 | 8611RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GA RCRA HINGED LID | MITG - Sharps Disposal | 002927 | 8611RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING ABSORBENT PADS | The Product 8611RC evaluated for the missing absorbent pads (absorbent pads/waste-toxic labels). The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: June 7, 2017 to Sep 10, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the kit item (absorbent pads/waste-toxic labels) packed prior to sealing the case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing kit item (absorbent pads/waste-toxic labels) for product should not be used and must be replaced. There is no patient involvement as it was discovered prior to use. So no safety or harm to end user. | 0702613906 | 0 | Y | United States | Closed | Customer reports: the case is missing all of the necessary labels, stickers, and absorbent pads. | 0 | 0 | 2018-06-07 | 2018-06-21 | 0 | 2018-06-07 | During use | 2018-09-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cardoso, Alexandra PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: CONTAINER - WASTE PHARMACY BLK-12 GAL Supplier #: KC8611RC Mfg. Part#: 8611RC Mfg: Covidien Lawson #: 577449 Looks like they are needing 10 of this set: (1) SharpSafety Absorbent Pad #R5097; (1) Set of Flammable Liquid 3 Red Labels - #R4370; (1) Set of Toxic Labels 6 White #R4371; and (8) Tamper Evident Seal Orange Labels #R4227 Below are the details for Product - Model No:8611RC, Description:12 GA RCRA HINGED LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 12 GA RCRA HINGED LID Model Number: 8611RC Serial or Lot Number: null ------------------------------------------------------------------------ ---------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Case is missing all of the necessary labels, stickers, and absorbent pads. Was the customer notified the product should be returned? No Please explain Customer does not need to return. Product still good. Just missing the labels and absorbent pads. PHYSICIAN LAST NAME: Moore PHYSICIAN FIRST NAME: Linda PHYSICIAN CLINIC: Memorial Hermann PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Humble PHYSICIAN STATE CODE: TX PHYSICIAN EMAIL: linda.moore@memorialhermann.org For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: Memorial Hermann FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 18951 Memorial North Drive FACILITY CITY: Humble FACILITY STATE CODE: TX FACILITY POSTAL CODE: 77338 FACILITY DAY PHONE: 281-319-8435 As Reported Event - departmentName: Diagnostic Imaging/Radiology SALESREP USER ID: cardoa8 SALESREP EMP ID: 285356 SALESREP NAME: Cardoso, Alexandra PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cardoso, Alexandra As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-19 | 0702617148-10 | 8881512852 | 17L26763X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO INFORMATION | AFC | KC:EXCESSIVE LUBRICANT | Note Creator: Daniel Callahan Note Date: 08/15/18 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. As of 08/15/18, no samples or pictures of the product that caused the customer complaint have been provided. A review of the device history record was performed for Lot# 17L26763X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881512852 was 172,800 parts (480 parts per case), produced on Line 23 on 12/03/17 – 12/04/17. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The most probable root cause is the spraymation heads were not cleaned daily during production of lot, which caused excessed silicone to be observed. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time. The corrective action to update the Daily Cleaning Log and will be tracked per Change Order EC113407. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0702617148 | 0 | Y | United States | Closed | Customer reported: Excessive globs of gummy oily substance inside the syringe. Way more than normal. Only one syringe that they seen. Hopefully that was the only one. They use these syringes for injections to patients and hope they haven't caused a problem. | PATTERSON VETERINARY HOSPITAL | 0012885050 | 2018-06-18 | 2018-06-19 | 0 | 2018-06-14 | Unknown | 2018-08-18 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-19 | 0702617180-10 | 8602RC | 18B01963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | (D) 2GA RCRA HAZARDOUSWASTE CONTX20 | MITG - Sharps Disposal | 002927 | 8602RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | Duplicate | #N/A | 0702617180 | 0 | Y | United States | Closed | Customer reported: Missing tampering labels on sharps containers | PATTERSON LOGISTICS SERVICES | 0012972160 | 2018-06-19 | 2018-06-19 | 0 | 2018-06-04 | Prior to Use | 2018-08-31 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-06-26 | 0702631737-10 | 65 | CL63R4422A | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 065 1 QUART PHLEBOTOMY CONTAINER | MITG - Sharps Disposal | 002927 | 65 |  |  |  | RFR | NEEDLE STICK/PUNCTURE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 8900SA evaluated for the contaminated needle stick. The reported condition cannot be observed as no sample/photo was available for product analysis. DHR Review was not performed as no lot number provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: June 25, 2017 to July 20, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The potential root cause associated with this event is improper handling/carrying of a filled sharps container for disposal by user. There was no information in the complaint regarding fill level for the container. As per the IFU carrying/lifting instructions “Never hold or carry used container below recommended fill line. Use additional caution with containers that do not have handles.” Also per IFU precaution “do not hold or carry container close to the body”. Also there is a caution on the label stating “Forcing could cause puncture”. The complaint has very limited information available to perform investigation as there is no photos/sample and lot number available. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The complaint has very limited information available. There was no information regarding lot number, fill level and any sample photos for further investigation. | 0702631737 | 0 | Y | United States | Closed | Customer Reported: a needle went through the sharps container and stuck me when I went to place it in our biohazard waste. | Meridian Health | 0005157014 | 2018-06-25 | 2018-06-26 | 0 | 2018-06-22 | During use | 2018-11-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Serious Illness Injury | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2018-06-27 | 0702635077-10 | 8881112599 | 17K27063X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NO INFORMATION | AFC | KC:MIXED COMPONENTS | Note creator: Daniel Callahan Note Date: 07/27/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. One (1) sample was received in a plastic bag directly from customer (Aquability). Luer and flange were damaged. Oral barrel received has print from Line 7 printer (cavity 45), but contains a plunger from Line 1 (cavity B47). A review of the device history record for Lot #17K27063X was performed with no discrepancies observed related to this defect. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112599 was 292,800 parts (800 parts per case), produced on Line 1 and on 10/19/17 – 10/22/17. 5 Why’s were used as the RCA investigative tool. Line 7 printed barrel was mixed with Line 1 barrels during transportation to Line 23 under special shop order. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time. Customer is able to perform 100% inspection for this type of barrel mixture, due to manufacturing process not designed for oral barrels. No other barrels from Line 7 were mixed. This complaint will be used for tracking and trending purposes. If ad ditional information is received warranting further analysis, the investigation may be resumed. | 0702635077 | 0 | Y | United States | Closed | The customer states:It is a very different syringe from the type we normally receive, in that it does not have a luer connection, only a smooth tip and it has printed gradations. It appears we have a product mix-up somewhere in the supply chain as we do not purchase any items such as this. | AQUABILITI | 0012728353 | 2018-06-20 | 2018-06-27 | 0 | 2018-06-01 | Prior to Use | 2018-08-24 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Daniel Callahan | 43308 | 43326 | Note Creator: Daniel Callahan Note Date: 07/27/18 Note Description: Addition of analysis One (1) sample was received in a plastic bag directly from customer (Aquability). Luer and flange were damaged. Oral barrel received has print from Line 7 printer (cavity 45), but contains a plunger from Line 1 (cavity B47). | Complete | #N/A
2018-06-29 | 0702641235-10 | 65 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 065 1 QUART PHLEBOTOMY CONTAINER | MITG - Sharps Disposal | 002927 | 65 |  |  |  | RFR | LOW PUNCTURE FORCE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 065 evaluated for the contaminated needle stick. The reported condition cannot be observed as no sample/photo was available for product analysis. DHR Review was not performed as no lot number provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: June 27, 2017 to July 20, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The probable root cause associated with this event may be improper handling/carrying of a filled sharps container for disposal below recommended fill line by user. There was no information in the complaint regarding fill level for the container. As per the IFU ca rrying/lifting instructions “Never hold or carry used container below recommended fill line. Use additional caution with containers that do not have handles.” Also per IFU precaution “do not hold or carry container close to the body”. Also there is a caution on the label stating “Forcing could cause puncture”. The complaint has very limited information available to perform investigation as there is no photos/sample and lot number available. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The complaint has very limited information available. There was no information regarding lot number, fill level and any sample photos for further investigation. | 0702641235 | 0 | Y | United States | Closed | Customer reports: a needle went through the sharps container giving the HCP a needle stick. | MERIDIAN WOMENS HEALTH CENTER | 0013065421 | 2018-06-27 | 2018-07-02 | 0 | 2018-06-22 | During use | 2018-11-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:065, Description:065 1 QUART PHLEBOTOMY CONTAINER IMPLANTABLE: No WHO HAS POSSESSION: Hospital REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 065 1 QUART PHLEBOTOMY CONTAINER Model Number: 065 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? Yes Provide details A needle went through the sharps container giving the HCP a needle stick Please provide a detailed description of the event A needle went through the sharps container and stuck the HCP when she went to place it in their biohazard waste. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Thornburg PHYSICIAN FIRST NAME: Denise PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Muncie PHYSICIAN STATE CODE: IN For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes FACILITY NAME: Meridian Health FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 100 N TILLOSTON AVE FACILITY CITY: MUNCIE FACILITY STATE CODE: IN FACILITY POSTAL CODE: 47304 SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-04 | 0702647167-10 | 8939 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The Product 8939 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: July 2, 2017 to Sep 12, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty four other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8939 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702647167 | 0 | Y | United States | Closed | The customer reported: The product was received with no lids in the case from the manufacturer. The case was sealed and had no damage. | St Anthony Hospital | 0015923170 | 2018-07-02 | 2018-07-04 | 0 | 0 | Prior to Use | 2018-09-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-05 | 0702649827-10 | 8989PG2 | 18B19163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8989PG2 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8989PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 18B19163 (3/18/18 to 3/19/18) and no internal reject found during production for the same issue. Trending Period: July 2, 2017 to July 25, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty four other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702649827 | 0 | Y | United States | Closed | The customer reports, the item was received without the lids. | Memorial Healthcare System | 0012710349 | 2018-07-02 | 2018-07-05 | 0 | 2018-05-17 | Unknown | 2018-07-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Jacqueline St Pierre | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-09 | 0702654702-10 | 8507SA | (10)18B11163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING CONTAINER | The Product 8507SA evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for lot 18B11163 (04/04/18 to 04/11/18) and no internal reject found during production for the reported condition. Trending Period: July 6, 2017 to Sep 14, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to follow standard work to pack out containers in the bottom layer and use weight scale for verification. The case quantity for 8507SA is 20 eaches/case. The packaging process is manual operation and weight scale is part of verification process. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing containers product should not be used. There was no patient involvement so no safety or harm to patient. | 0702654702 | 0 | Y | United States | Closed | The customer reported one case had 20 lids but only 10 containers. | GIBSON COMMUNITY HOSPITAL | 0013256451 | 2018-07-09 | 2018-07-09 | 0 | 2018-07-06 | Prior to Use | 2018-09-25 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Jacqueline St Pierre | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-10 | 0702657550-10 | 8939 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8939 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: July 10, 2017 to July 25, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty four other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The complaint is from outside distribution. The case quantity for 8939 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702657550 | 0 | Y | United States | Closed | Customer received 1cs of the item without lids. | UNIVERSITY HEALTH CARE SYSTEM | 0013239456 | 2018-07-10 | 2018-07-10 | 0 | 2018-07-03 | Unknown | 2018-11-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-11 | 0702660843-10 | 65 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 065 1 QUART PHLEBOTOMY CONTAINER | MITG - Sharps Disposal | 002927 | 65 |  |  |  | RFR | NEEDLE STICK/PUNCTURE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 065 evaluated for the contaminated needle stick. The reported condition cannot be observed as no sample/photo was available for product analysis. DHR Review was not performed as no lot number provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: June 26, 2017 to Aug 16, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaint reported for the product family. Root Cause: The probable root cause associated with this event may be improper handling/carrying of a filled sharps container for disposal below recommended fill line by user. There was no information in the complaint regarding fill level for the container. As per the IFU ca rrying/lifting instructions “Never hold or carry used container below recommended fill line. Use additional caution with containers that do not have handles.” Also per IFU precaution “do not hold or carry container close to the body”. Also there is a caution on the label stating “Forcing could cause puncture”. The complaint has very limited information available to perform investigation as there is no photos/sample and lot number available. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event for the product code 065. The complaint has very limited information available. There was no information regarding lot number, fill level and any sample photos for further investigation. | 0702660843 | 0 | Y | United States | Closed | The customer states a needle went through the sharps container and stuck the clinician when she was placing it in the biohazard bin.-###-From PE-702641235-Customer reports: a needle went through the sharps container giving the HCP a needle stick. | MERIDIAN WOMENS HEALTH CENTER | 0013065421 | 2018-06-26 | 2018-07-16 | 0 | 2018-06-22 | During use | 2019-01-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Copeland, Annie PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Other COMPLAINT SOURCE OTHER: McKesson ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:065, Description:065 1 QUART PHLEBOTOMY CONTAINER IMPLANTABLE: No WHO HAS POSSESSION: Hospital REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 065 1 QUART PHLEBOTOMY CONTAINER Model Number: 065 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Unknown, but will follow-up for more information Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event A needle went through the sharps container and stuck the clinician when she was placing it in the biohazard bin Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Thornburg PHYSICIAN FIRST NAME: Denise PHYSICIAN CLINIC: Meridian Health PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 100 N Tillotson Ave PHYSICIAN CITY: Muncie PHYSICIAN STATE CODE: IN PHYSICIAN POSTAL CODE: 47304 PHYSICIAN DAY PHONE: 765-744-7577 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: copela3 SALESREP EMP ID: 385088 SALESREP NAME: Copeland, Annie PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Bradley, Nicholas ADDITIONAL COMMENTS ABOUT THE EVENT: I have left 4 messages for Denise but have not heard back As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes -###-From PE-702641235- FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:065, Description:065 1 QUART PHLEBOTOMY CONTAINER IMPLANTABLE: No WHO HAS POSSESSION: Hospital REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 065 1 QUART PHLEBOTOMY CONTAINER Model Number: 065 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? Yes Provide details A needle went through the sharps container giving the HCP a needle stick Please provide a detailed description of the event A needle went through the sharps container and stuck the HCP when she went to place it in their biohazard waste. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Thornburg PHYSICIAN FIRST NAME: Denise PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Muncie PHYSICIAN STATE CODE: IN For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes FACILITY NAME: Meridian Health FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 100 N TILLOSTON AVE FACILITY CITY: MUNCIE FACILITY STATE CODE: IN FACILITY POSTAL CODE: 47304 SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2018-07-11 | 0702661054-10 | 8989 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: July 11, 2017 to July 27, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty three other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702661054 | 0 | Y | United States | Closed | The customer reports missing lids upon arrival. | DEKALB MEDICAL CENTER | 0012597590 | 2018-07-11 | 2018-07-12 | 0 | 2018-07-11 | During use | 2018-10-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Maxfield, Peter PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Chemo Protection Safety --------------------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989, Description:8989 18 GAL CONT CHEMO HNGD LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8989 18 GAL CONT CHEMO HNGD LID X5 Model Number: 8989 Serial or Lot Number: null ------------------------------------------------------------------------ --------------------------- When was the issue noticed? Before use/during setup Please provide a detailed description of the event Missing lids upon arrival. For Physician with Siebel ID 0015654870: PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: maxfip1 SALESREP EMP ID: 315121 SALESREP NAME: Maxfield, Peter PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Maxfield, Peter As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0015654870 Debbie Perry cannot find a BP in GCH | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-18 | 0702676166-10 | 8982 | 171782 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8982 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 171782 (08/02/16 to 08/05/16) and no internal reject found during production for the same issue. Trending Period: July 18, 2017 to Aug 9, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702676166 | 0 | Y | United States | Closed | The customer reports case arrived with 4 missing lids. Containers are still in tact but lids need to be replaced. | REDLANDS COMMUNITY HOSPITAL | 0001109113 | 2018-07-18 | 2018-07-24 | 0 | 2018-07-18 | During use | 2018-08-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Daniel, Alan-Anthony PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Chemo Protection Safety --------------------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Other COMPLAINT SOURCE OTHER: Alicia Sanchez ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8982, Serial/Lot No:171782, Description:8982 2GAL YEL CONTAINER CHEMO X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8982 2GAL YEL CONTAINER CHEMO X20 Model Number: 8982 Serial or Lot Number: 171782 ------------------------------------------------------------------------ ---------------------------- When was the issue noticed? Before use/during setup Please provide a detailed description of the event Case arrived with 4 missing lids. Containers are still in tact but lids need to be replaced. SALESREP USER ID: daniea17 SALESREP EMP ID: 349087 SALESREP NAME: Daniel, Alan-Anthony PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daniel, Alan-Anthony As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-19 | 0702678268-10 | 8989 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: July 19, 2017 to Aug 9, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twenty three other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. The packaging process is manual operation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702678268 | 0 | Y | United States | Closed | The customer received items without lids. | UNIVERSITY HEALTHCARE ALLIANCE | 0012863579 | 2018-07-16 | 2018-07-19 | 0 | 2018-06-19 | Unknown | 2018-11-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-20 | 0702681382-10 | 8989CRH | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 18 GAL CHEMO RED HINGE LID | MITG - Sharps Disposal | 002927 | 8989CRH |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | Duplicate | #N/A | 0702681382 | 0 | N | Canada | Closed | The customer states that the red large sharps containers have lids that are very difficult to snap onto the bottom of the container. There are times when they will not snap safely into place. | YORKTON REGIONAL HEALTH CENTRE | 0012629903 | 2018-07-20 | 2018-07-23 | 0 | 2018-04-27 | During use | 2018-08-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Danos, Maria PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Not specified Did the event lead to or extend patient hospitalization? Not specified Was there any injury as a result of the event? Not specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Are requesting we re-design these to make them more user friendly. No issues with the yellow containers. Below are the details for Product - Model No:8989CRH, Description:18 GAL CHEMO RED HINGE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 18 GAL CHEMO RED HINGE LID Model Number: 8989CRH Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Customer is complaining that the red large sharp containers have lids that are very difficult to snap onto the bottom of the container. There are times that they will not snap safely into place. PHYSICIAN LAST NAME: Fahlman PHYSICIAN FIRST NAME: Jeanette PHYSICIAN CLINIC: Yorkton Regional Health Centre PHYSICIAN COUNTRY CODE: CAN PHYSICIAN ADDRESS1: 270 Bradbrooke Dr. PHYSICIAN CITY: Yorkton PHYSICIAN STATE CODE: SK PHYSICIAN DAY PHONE: (306)786-0448 For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes SALESREP USER ID: danosm1 SALESREP EMP ID: 314177 SALESREP NAME: Danos, Maria PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Danos, Maria As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-22 | 0702683055-10 | 8860 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGTMPDIS | 12GAL PHARMSAFETY WASTE CNT 10 | MITG - Temp Mgmt Fluid Warmer | 003173 | 8860 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8860 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: July 21, 2017 to Aug 10, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and ten other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8860 is 10 eaches/case. The packaging process is manual operation. There was no lot number available to perform DHR review. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement so no safety or harm to patient. | 0702683055 | 0 | Y | United States | Closed | Customer reported: Missing the lids. | BERKLEY EAST CONV HOSPITAL | 0013273411 | 2018-07-21 | 2018-07-22 | 0 | 0 | Prior to Use | 2018-11-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-07-26 | 0702692607-10 | 8934CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8934CRS |  |  |  | RFR | CRACKED CONTAINER | AFC | Duplicate | #N/A | 0702692607 | 0 | N | Canada | Closed | The customer reports the base of the bin cracked causing a chemo spill. | FRASER VALLEY CANCER CENTRE | 0012689372 | 2018-07-25 | 2018-07-27 | 0 | 2018-06-13 | During use | 2018-08-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | #N/A
2018-09-05 | 0702760585-10 | UNK SHARP CONT | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | UNKNOWN SHARP CONTAINER | MITG - Sharps Disposal | 002927 | UNK SHARP CONT |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING LIDS | The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review: Review of device history record could not be performed. Trending Period: N/A. Product Code unknown. Trending Results: N/A Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity is unknown, as code is unknown. The packaging process is manual operation. Formal Investigation: Based on the lack of information available, no formal investigation will be required. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702760585 | 0 | Y | United States | Closed | Customer reports: the box arrived unsealed and missing a container lid. | FRESENIUS MEDICAL CARE NA | 0012833563 | 2018-08-31 | 2018-09-13 | 0 | 0 | Prior to Use | 2018-10-04 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-09-06 | 0702761877-10 | 8935Y | 17J03663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8935Y YEL 12GAL SHARP CONT SLIDING LID | MITG - Sharps Disposal | 002927 | 8935Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:BROKEN CONTAINER | Product Analysis: As of 10/04/18 no physical samples received. Three (3) images were attached to complaint and analyzed. DHR Review: Review of Device History Record showed no discrepancies found relating to defect reported. Trending: A review of customer complaints as well as internal NCR rejects pertaining to “Cracked Containers” for this product family was performed. Two (2) complaints for “Cracked Container” were received in the trending period from 09.05.17 – 10.04.18. Zero (0) internal NCR rejects of this nature for this product family were received in the trending period from 07.01.18 – 10.04.18. Root Cause: The most probable root cause is the containers may have cracked during shipping and handling due external stress/forces on the container. Formal Investigation: Based on the information available, trend analysis, and the investigation findings, a CAPA decision tool will be completed. Complaint will be re-investigated if additional information is supplied from customer. | 0702761877 | 0 | Y | United Arab Emirates | Closed | The customer states the containers were found cracked when the carton was opened. | NMC | 0011072138 | 2018-09-05 | 2018-09-10 | 0 | 0 | Prior to Use | 2018-12-05 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2018-09-19 | 0702782759-10 | 8938 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:WARPED LIDS | Product Analysis: As of 10/12/18, no physical samples or pictures were received by customer. Product analysis cannot be performed until the latter is received. DHR Review: Review of Device History Record could not be performed. Trending: A review of customer complaints as well as internal NCR rejects pertaining to “warped lids” for this product family was pe rformed. Zero (0) complaints for “warped lids” were received in the trending period from 09.17.17 – 10.12.18. Zero (0) internal NCR rejects of this nature for this product family were received in the trending period from 06.17.18 – 10.12.18. Root Cause: The complaint event description states that “part of the lid shape was changed.” The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces which may cause lid warpage. Formal Investigation: Based on the information available, trend analysis, and the investigation findings, no formal investigation is required. Complaint will be re-investigated if additional information is supplied from customer. | 0702782759 | 0 | Y | Japan | Closed | 0 | Kobe City Medical Center | 0009364489 | 2018-09-18 | 2018-09-19 | 0 | 2018-09-17 | Prior to Use | 2018-10-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-10-09 | 0702813917-10 | 8982 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MALFUNCTIONING LID | Investigation Date: 11/16/18 Investigator: Daniel Callahan Note: Summary of Investigation Product Analysis: One (1) sample was received from customer and analyzed on 11/07/18. Yellow lid was detached from lid when received. DHR Review: A review of Lot #18B23963 was performed with no discrepancies found that are related to the failure mode reported. All inspections were completed and met quality acceptance criteria per established sampling plans. Lids with yellow caps were pulled from Lots 181353 and 180996. These were produced on 02/10/18 and 12/27/17 respectively. Trend Analysis: A review of customer complaints as well as internal NCR rejects pertaining to this failure mode and product family was performed. Zero (0) complaints for “malfunctioning lid” related to the yellow cap were received in the trending period from 10.04.17 – 11.10.18. Zero (0) internal NCR rejects of this nature for this product family were received in the trending period from 08.10.18 – 11.10.18. Root Cause: The possible root cause of this failure was rotary actuator malfunction. Turns the head of the robot that picks up yellow cap and places into correct position on lid. Formal Investigation: Based on the information available, trend analysis, and the investigation findings, no formal investigation will be initiated. AM-2 05 assembly machine was inspected and verified to be welding yellow caps to lids correctly. Reported defect is determined to be an isolated incident. | 0702813917 | 0 | Y | Japan | Closed | Customer reports: The yellow lid wasn't fixated. | Niti-on Co., Ltd | 0015877231 | 2018-10-04 | 2018-10-09 | 0 | 2018-10-03 | Prior to Use | 2018-11-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2018-10-09 | 0702813918-10 | 8938 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:NO FAILURE | Date: 11/16/18 Investigator: Daniel Callahan Note: Summary of Investigation Product Analysis: One (1) sample was received from customer and analyzed on 11/07/18. Lid was received with no container. Lid was in “Final Lock” position and could not be opened. DHR Review: A review of the DHR could not be performed. Lot # not reported. Trend Analysis: A review of customer complaints as well as internal NCR rejects pertaining to this failure mode and product family was performed. Zero (0) complaints for “malfunctioning lid” related to the yellow cap were received in the trending period from 10.04.17 – 11.10.18. Zero (0) internal NCR rejects of this nature for this product family were received in the trending period from 08.10.18 – 11.10.18. Root Cause: N/A. Customer reported failure was not confirmed. Lid was locked and could not be opened. Formal Investigation: Based on the information available, trend analysis, and the investigation findings, no formal investigation will be initiated. Repo rted failure could not be confirmed. Sample received was locked and could not be opened, contrary to customer report. | 0702813918 | 0 | Y | Japan | Closed | Customer reports: the locked lid opened up. | Kobe City Medical Center West Hospital | 0011150506 | 2018-10-04 | 2018-10-09 | 0 | 2018-10-03 | Prior to Use | 2018-11-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | Daniel Callahan | 43420 | 43420 | One (1) sample was received from customer and analyzed on 11/07/18. Lid was received with no container. Lid was in “Final Lock” position and could not be opened. | Complete | Yes
2018-12-18 | 0702923615-10 | 8980S | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980S 8GAL RED CONT BIOHAZ SLIDETOPX10 | MITG - Sharps Disposal | 002927 | 8980S |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | KC:MALFUNCTIONING LID | Product Analysis: One (1) lid sample was received and analyzed on 01/03/19. Sample received had 8G sliding top (clear) reversed. Side panel is cracked, where customer attempted to correct orientation as described in notes attached. DHR Review: No DHR review performed. Lot # not reported. Root Cause: Root Cause: Operator miss-assembled sliding lid in reverse position. Trending Analysis: A review of customer complaints as well as internal NCR rejects for “Malfunctioning Lid” failure modes and product family was performed. Zero other incidents of this nature for this product family was received in the trending period from 01.28.18 – 01.28.19. Zero (0) internal NCR rejects of this nature for this product family were received in the trending period from 10.28.18 – 01.28.19. Conclusion: Based on the information available, trend analysis, and the investigation findings, a new formal investigation is not deemed necessa ry at this time. This complaint will be used for tracking and trending purposes. | 0702923615 | 0 | Y | Japan | Closed | Customer reports: The sliding lid was placed in the opposite position. After additional information, the customer reports that the lid can be reversed but will not lock. | Kawasaki Medical School Hospital | 0008034829 | 2018-12-17 | 2018-12-18 | 0 | 2018-12-16 | Prior to Use | 2019-02-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Daniel Callahan | 43503 | 43503 | Investigator: Daniel Callahan Date: 01/28/19 One (1) lid sample was received and analyzed on 01/03/19. Sample received had 8G sliding top (clear) reversed. Side panel is cracked, where customer attempted to correct orientation as described in notes attached. | Complete | Yes
2019-01-09 | 0702949250-10 | 8931 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8931 12GAL CONT CHEMO W/ WHI HNGD LIDX10 | MITG - Chemo Protection Safety | 003209 | 8931 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MISSING ABSORBENT PADS | Note Creator: Daniel Callahan Date: 02/13/19 Title: Investigation Summary Product Analysis: One (1) box was received with absorbent pad missing. Ten (10) lids and containers were present with only 1 bag of absorbent pads included. DHR Review: DHR review was performed. Records show Lot was produced per required specifications and met quality acceptance criteria. Root Cause: Operator failed to include 2nd absorbent pad in container during packaging. Trending Analysis: A review of customer complaints as well as internal NCR rejects for “Missing absorbent pads” failure mode and product family was performed. One (1) other incident of this nature for this product family was received in the trending period from 01.17.18 – 02.13.19. One (1) internal NCR rejects of this nature for this product family were received in the trending period from 11.13.18 – 02.13.19. Conclusion: Based on the information available, trend analysis, and the investigation findings, a new formal investigation is not deemed necessa ry at this time. Failure is low risk. This complaint will be used for tracking and trending purposes. | 0702949250 | 0 | Y | Hong Kong | Closed | The customer states: 1 bag of absorbent pad was missing from the box. | 0 | 0 | 2018-12-20 | 2019-01-17 | 0 | 2018-11-05 | Prior to Use | 2019-02-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Daniel Callahan | 43509 | 43510 | One (1) box was received with absorbent pad missing. Ten (10) lids and containers were present with only 1 bag of absorbent pads included. | Complete | #N/A
2017-02-24 | 0701684622-10 | 8506Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8506Y YEL 5 QRT CONT S STAR C B DR X20 | MITG - Sharps Disposal | 002927 | 8506Y |  |  |  | RFR | LID ISSUE - FAILURE TO DROP SHARP | AFC | KC:MALFUNCTIONING LID | The Product 8506Y evaluated for the failure to drop sharp-malfunction lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Feb 20, 2016 to April 17, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: As per the complaint description container was free standing. For container stabilization Medtronic hardware should be used. As per the complaint description no hardware used with the device. It is recommended that 8506Y product shall be used with hardware for stability during use. Also as per the complaint description it seems like the product being used is not suitable for the end use application. It may be different style of lid-horizontal drop opening lid shall be used in this type of setting. As per additional information from sales rep different style lid opening product was suggested as alternate for customer’s specific application and also discussed about holder for sharp containers for stability. The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be used with hardware for sharp container’s stability. Also as per sales rep note different style lid opening unit should be used for customer’s specific end use application. | 0701684622 | 0 | Y | Canada | Closed | The customer reports that sharps are getting stuck on the flap. When they are very busy they don't notice that sharps have not dropped and in one case, the container that was on the ledge tipped over and needles spilled out onto the floor creating safety risk. There was no patient involvement. | Eagle Ridge Hospital | 0008053914 | 2017-02-20 | 2017-03-09 | 0 | 2017-02-20 | Post-Operative | 2017-11-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury - Potential for Injury | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-02-28 | 0701690339-10 | 8970Y | 00172667 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | CRACKED CONTAINER | AFC | KC:BROKEN CONTAINER | The Product 8970Y evaluated for the cracked container. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met the specification. DHR Review: Device History Records data reviewed for 172667 (09/25/2016 to 09/28/2016). The internal NCR 16090028 was issued for missing inserts into containers per standard work. The issue was identified at start up and all affected product was contained as part of NCR. The product was reworked and inspected per specified AQL and passed the AQL inspection. Trending Period: Feb 22, 2016 to April 17, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed one internal inspection reject and no other complaints reported for the product code. Root Cause: The DHR was reviewed and it was found that it had NCR16090028 for missing inserts per standard work. The issue was identified at start up and all affected product was contained as part of NCR16090028. The product was reworked, inspected per specified AQL and passed AQL. The most probable root cause is human error- failure to put insert prior to seal the case and as NCR rework is inspected per AQL. Also outside of Medtronic, other distribution center may have repackaged product. Medtronic has Case Quantity shipped 20 Eaches per case for 8970Y. As per the complaint no. of defective quantity is 1 each, which shows that case may had been broken down to small quantity by distribution center to ship to end user. Also there is no evidence documented in complaint about container is inspected for any damage prior to use. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. Also there is no information about if customer received full case or only one each for use. | 0701690339 | 0 | Y | Canada | Closed | The customer states the bottom of the sharps container broke when the nurse was preparing to remove the container. Bottom broke out and contents spilled out, clients in room nurse moved them out of the room so that there was no danger to them. Container contained dirty sharps so disposed of as per regular process. | Northwest Community Health Center | 0011800149 | 2017-02-27 | 2017-03-20 | 0 | 2017-02-22 | During use | 2017-08-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | Yes
2017-03-02 | 0701695097-10 | 8881512878 | 16K04863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | CONTAMINATION | AFC | KC:DAMAGED COMPONENTS | Incident Description: The customer reports staff member opened a syringe and found her hands slimy. Looked at the syringe and realized some sort of clear substance was on the outside and inside of the syringe. Also realized that end of syringe barrel broken, which was what caused the plunger to come out. Huge potential for harm depending on what the substance is, as this syringe was going to be hooked up to a central line. Product was kept. Syringe did not make it to the patient, it was caught prior. Product Analysis: One sample was received from the customer. The sample arrived in two pieces; (1) printed barrel and (2) rigid pack. (1) An inspection of the barrel was performed under normal vision. The reported condition of “broken barrel” was observed, however the reported condition of “clear substance” was not readily apparent under normal vision. The barrel was then placed under lighted magnification (3 diopter). The inspection confirmed a clear substance inside the barrel with trace amounts on outside within 3/8” to the top. It was also noted that the substance was evenly dispersed. Droplets/particles were not observed. - DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512878 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. (2) The tamper evident seal on the rigid pack cap was broken. Upon removal of the cap, the assembled plunger rod was found free floating in the sleeve. DHR Review: A review of the device history record for lot 16K04863X was completed. 360 cases (172,800 pieces) were released on 11/16/16 using production line 23. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M assessment that addressed: Man, Mat erial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed as a corrective and preventative action where appropriate. Silicone application is part of the manufacturing process and a required component of the bill of material. In the event of a broken barrel, silicone may be applied to the top outside portion of the barrel. As such, the 6M assessment focuses on broken barrel. Man Causes: Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Assembly System Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Material Causes: Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met). Method Causes: Insufficient Standard Operating Procedure (SOP) for Line 2/3 Syringe Assembly Machine (Not Significant – Method is sufficient based on trending results). Insufficient Standard Operating Procedure (SOP) for setting/troubleshooting the Assembly Machine barrel drop station (Significant – Root Cause). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Machine Causes: Assembly machine barrel drop station not setup properly (Significant) Process parameters not maintained (Not significant – Parameters are verified throughout production lots) Aidlin parts jamming (Significant) Printer parts jamming (Not Significant – Part(s) that jam during printing operation are not likely to make it all the way through the automated assembly system). Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Damage Functional Less Than Desired) and product family was performed. Zero other complaints of this nature for this product family were received in the trending period from 2.24.16 – 03.31.17. Zero internal rejects of this nature for this product family were received in the trending period from 12.31.16 – 03.31.17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. The below improvement actions were previously implemented on 01/06/17. Furthermore, zero other related non-conformances have been identified during trending. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: DC00004509 Clearing Jams on Aidlins and DC00058484 Troubleshooting the barrel drop station on the RAM were created and released to production on 01/06/17. These changes ensure proper handling of aidlin jams and setup/troubleshooting of assembly machine barrel drop stations. The newly created operating procedures will prevent future instance of damage being created. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure and root cause for “Damage Functional Less Than Desired” is reflected. | 0701695097 | 0 | Y | Canada | Closed | The customer reports staff member opened a syringe and found her hands slimy. Looked at the syringe and realized some sort of clear substance was on the outside and inside of the syringe. Also realized that end of syringe barrel broken, which was what caused the plunger to come out. Huge potential for harm depending on what the substance is, as this syringe was going to be hooked up to a central line. Product was kept. Syringe did not make it to the patient, it was caught prior. | St Pauls Hospital | 0016127126 | 2017-02-27 | 2017-03-21 | 0 | 2017-02-24 | During use | 2017-09-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury - Potential for Injury | FIRST NOTIFIED NAME: Sparks, James richard ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512878, Serial/Lot No:16K04863X, Description:8881512878 12CC S-LK SYR ONLY O/P X480 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881512878 12CC S-LK SYR ONLY O/P X480 Model Number: 8881512878 Serial or Lot Number: 16K04863X ------------------------------------------------------------------------ ------------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Unknown, but will follow-up for more information Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Larsen PHYSICIAN FIRST NAME: Jennifer PHYSICIAN CLINIC: St. Pauls Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Saskatoon PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Cote Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 FACILITY NAME: St. Pauls Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: Kidney Health SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Shanika Hodge | Jeremy Roberts | 42825 | 42825 | Product Analysis: One sample was received from the customer. The sample arrived in two pieces; (1) printed barrel and (2) rigid pack. (1) An inspection of the barrel was performed under normal vision. The reported condition of “broken barrel” was observed, however the reported condition of “clear substance” was not readily apparent under normal vision. The barrel was then placed under lighted magnification (3 diopter). The inspection confirmed a clear substance inside the barrel with trace amounts on outside within 3/8” to the top. It was also noted that the substance was evenly dispersed. Droplets/particles were not observed. - DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512878 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. (2) The tamper evident seal on the rigid pack cap was broken. Upon removal of the cap, the assembled plunger rod was found free floating in the sleeve. | Complete | Yes
2017-03-02 | 0701695117-10 | 8881512852 | 16G06063X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | US:LUBRICATION EXCESSIVE | Incident Description: The customer states the syringe has an unknown slimy/greasy substance on the rubber stopper. Product Analysis: Forty samples were returned for review. The samples arrived in two separate bags labeled by product event number and line item number, which directly correlate with the lot numbers reported. 1 - 701695117-10 (lot 16G06063X) and 2 - 701695117-20 (lot 16G06163X). 1 – 23 samples were provided. 21 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 2 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. 2 – 17 samples were provided. 6 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 11 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. All samples were sent to the Medtronic Technical Services Chemistry Lab for evaluation against the ISO 7886-1 standard to determine if lubricate levels were within the established ISO acceptable limit. The report confirmed that all samples were well within the limit with the highest sample only reading at 39% of the limit. DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512852 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Excess silicone is considered as a Minor defect in accordance with established quality and inspection standards. DHR Review: A review of the device history record for lots 16G06063X & 16G06163X was completed. 602 cases (288,960 pieces) were released on 10/25/16 using production line 23. The DHR was reviewed and no abnormal process conditions were present during the manufacturing of this product that could have contributed to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: -Spraymation head levels were set incorrectly (Not significant – parameters were verified at each start up) -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective) Material Causes: -Use of incorrect silicone (Not significant – incoming inspection processes are in place and being followed) Method Causes: -Insufficient procedure for RAM cleaning (Significant – Cleaning/Wiping down of spraymation heads is not included daily cleaning activities. The heads have a potential of building up silicone during operation, with the potential of small drops falling into assembled syringes). -Insufficient preventative maintenance for Spraymation heads (Not Significant – A Preventative Maintenance plan is in place) -Process not validated (Not significant – spraymation process was confirmed to be validated) Measurement Causes: -QA inspection for spraymation testing not performed (Not significant – the inspection was confirmed to be performed as required) -Process parameters at start up were not verified (Not significant – parameters were verified) Machine Causes: -Spraymation heads not functioning properly/malfunction (Not Significant – The heads are confirmed to be on a regular rebuild schedule/PM) -Silicone delivery system not operating properly (Not significant – operation is confirmed via Syringe actuation testing throughout the lots) Environment Causes: -Environmental controls were out of validated operating ranges (Not significant – environmental controls are observed and documented once per shift) Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Excessive Lubricate in Syringe) and product family was performed. 3 other complaints of this nature for this product family were received in the trending period from 5.10.17 – 03.02.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.10.17 – 02.2.17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC023506 has been initiated to update daily cleaning activities, requiring cleaning/wiping down of spraymation heads. This cleaning will prevent build-up of silicone during manufacturing, decreasing the likelihood of silicone from falling into assembled syringes Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of excess silicone. Cause is due to normal manufacturing. Current process controls references Cleanliness GS20002641, which is the procedure which was updated. | 0701695117 | 0 | Y | Canada | Closed | The customer states the syringe has an unknown slimy/greasy substance on the rubber stopper. | ST PAULS HOSPITAL | 0012563337 | 2017-03-02 | 2017-03-21 | 0 | 2017-03-02 | During use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Sparks, James richard ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512852, Serial/Lot No:16G06063X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, replaced by another manufacturer's product PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 16G06063X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes Below are the details for Product - Model No:8881512852, Serial/Lot No:16G06163X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, replaced by another manufacturer's product PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 16G06163X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Larsen PHYSICIAN FIRST NAME: Jennifer PHYSICIAN CLINIC: St. Pauls Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Saskatoon PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Cote Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 FACILITY NAME: St. Pauls Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: Kidney Health SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes Removed from Brief Description: St. Pauls 12ML Syringe March 2 2107 | Jacqueline St Pierre | Jeremy Roberts | 42865 | 42867 | Product Analysis: Forty samples were returned for review. The samples arrived in two separate bags labeled by product event number and line item number, which directly correlate with the lot numbers reported. 1 - 701695117-10 (lot 16G06063X) and 2 - 701695117-20 (lot 16G06163X). 1 – 23 samples were provided. 21 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 2 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. 2 – 17 samples were provided. 6 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 11 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. All samples were sent to the Medtronic Technical Services Chemistry Lab for evaluation against the ISO 7886-1 standard to determine if lubricate levels were within the established ISO acceptable limit. The report confirmed that all samples were well within the limit with the highest sample only reading at 39% of the limit. DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512852 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Excess silicone is considered as a Minor defect in accordance with established quality and inspection standards. | Complete | #N/A
2017-03-02 | 0701695117-20 | 8881512852 | 16G06163X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | US:LUBRICATION EXCESSIVE | Incident Description: The customer states the syringe has an unknown slimy/greasy substance on the rubber stopper. Product Analysis: Forty samples were returned for review. The samples arrived in two separate bags labeled by product event number and line item number, which directly correlate with the lot numbers reported. 1 - 701695117-10 (lot 16G06063X) and 2 - 701695117-20 (lot 16G06163X). 1 – 23 samples were provided. 21 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 2 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. 2 – 17 samples were provided. 6 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 11 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. All samples were sent to the Medtronic Technical Services Chemistry Lab for evaluation against the ISO 7886-1 standard to determine if lubricate levels were within the established ISO acceptable limit. The report confirmed that all samples were well within the limit with the highest sample only reading at 39% of the limit. DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512852 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Excess silicone is considered as a Minor defect in accordance with established quality and inspection standards. DHR Review: A review of the device history record for lots 16G06063X & 16G06163X was completed. 602 cases (288,960 pieces) were released on 10/25/16 using production line 23. The DHR was reviewed and no abnormal process conditions were present during the manufacturing of this product that could have contributed to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: -Spraymation head levels were set incorrectly (Not significant – parameters were verified at each start up) -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective) Material Causes: -Use of incorrect silicone (Not significant – incoming inspection processes are in place and being followed) Method Causes: -Insufficient procedure for RAM cleaning (Significant – Cleaning/Wiping down of spraymation heads is not included daily cleaning activities. The heads have a potential of building up silicone during operation, with the potential of small drops falling into assembled syringes). -Insufficient preventative maintenance for Spraymation heads (Not Significant – A Preventative Maintenance plan is in place) -Process not validated (Not significant – spraymation process was confirmed to be validated) Measurement Causes: -QA inspection for spraymation testing not performed (Not significant – the inspection was confirmed to be performed as required) -Process parameters at start up were not verified (Not significant – parameters were verified) Machine Causes: -Spraymation heads not functioning properly/malfunction (Not Significant – The heads are confirmed to be on a regular rebuild schedule/PM) -Silicone delivery system not operating properly (Not significant – operation is confirmed via Syringe actuation testing throughout the lots) Environment Causes: -Environmental controls were out of validated operating ranges (Not significant – environmental controls are observed and documented once per shift) Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Excessive Lubricate in Syringe) and product family was performed. 3 other complaints of this nature for this product family were received in the trending period from 5.10.17 – 03.02.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.10.17 – 02.2.17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC023506 has been initiated to update daily cleaning activities, requiring cleaning/wiping down of spraymation heads. This cleaning will prevent build-up of silicone during manufacturing, decreasing the likelihood of silicone from falling into assembled syringes Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of excess silicone. Cause is due to normal manufacturing. Current process controls references Cleanliness GS20002641, which is the procedure which was updated. | 0701695117 | 0 | Y | Canada | Closed | The customer states the syringe has an unknown slimy/greasy substance on the rubber stopper. | ST PAULS HOSPITAL | 0012563337 | 2017-03-02 | 2017-03-21 | 0 | 2017-03-02 | During use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Sparks, James richard ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512852, Serial/Lot No:16G06063X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, replaced by another manufacturer's product PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 16G06063X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes Below are the details for Product - Model No:8881512852, Serial/Lot No:16G06163X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, replaced by another manufacturer's product PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 16G06163X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Larsen PHYSICIAN FIRST NAME: Jennifer PHYSICIAN CLINIC: St. Pauls Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Saskatoon PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Cote Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 FACILITY NAME: St. Pauls Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: Kidney Health SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes Removed from Brief Description: St. Pauls 12ML Syringe March 2 2107 | Jacqueline St Pierre | Jeremy Roberts | 42865 | 42867 | Product Analysis: Forty samples were returned for review. The samples arrived in two separate bags labeled by product event number and line item number, which directly correlate with the lot numbers reported. 1 - 701695117-10 (lot 16G06063X) and 2 - 701695117-20 (lot 16G06163X). 1 – 23 samples were provided. 21 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 2 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. 2 – 17 samples were provided. 6 of the samples remained sealed in the original primary package preventing review of the reported condition without opening (breaking the tamper evident seal). 11 samples were provided unsealed. All samples were visually inspected in the fully assembled and provided state under normal vision. A slimy/greasy substance on the rubber stopper was not observed. Each syringe was then actuated by pulling back on the plunger. A small silicone drop was observed on few samples. The plunger rod was then fully disassembled and removed from the syringe barrel. Upon removal, the reported condition was further observed. All samples were sent to the Medtronic Technical Services Chemistry Lab for evaluation against the ISO 7886-1 standard to determine if lubricate levels were within the established ISO acceptable limit. The report confirmed that all samples were well within the limit with the highest sample only reading at 39% of the limit. DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512852 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Excess silicone is considered as a Minor defect in accordance with established quality and inspection standards. | Complete | #N/A
2017-03-03 | 0701697203-10 | 8881112599 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:MISSING COMPONENTS | Incident Description: Customer reports that one bag of syringes lot # 17A02563X had incomplete assembly of syringes - plungers were missing. Product Analysis: One box/bag of 800 samples was returned for evaluation. The original product identification/shipping label was covered with a white label and thus lot number was unable to be confirmed. The box was intact and shipping damage was not observed. The package was opened and all 800 samples were visually inspected. Of 800 samples returned for evaluation, 159 were found to be missing a plunger rod and rubber tip (Incomplete assembled syringes). In the return bag sent by Aquabiliti the missing plungers could be observed as reported. DHR Review: A review of the device history record for lot 17A02563X was completed. 245 cases were completed on 01/14/2017 using production line 1. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M assessment that addressed: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed as a corrective and preventative action when appropriate. Man Causes: -Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records). -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective). Material Causes: - Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met). - Rubber Tips (Not significant – An analysis is performed on tips at point of manufacturing to ensure all characteristics meet performance requirements. Incoming inspection is performed upon delivery of the tips to ensure requirements are met). Method Causes: Insufficient Standard Operating Procedure (Significant: SOP does not clearly define requirements of putting the assembly machine back in service). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Machine Causes: Assembly machine plunger drop station not monitored properly (Significant) Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Based on the investigation, the most probable root cause is that spraymation testing was performed in accordance with LS20001213 during change over from lot 16M04563X to the reported lot and the assembly machine was not confirmed to be put back in service upon completion. This allowed assemblies to be produced that were not fully assembled. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Missing Components) and product family was performed. Zero other complaints of this nature for this product family were received in the period 2/27/16 to 5/24/17. Zero internal rejects of this nature for this product family were received in the period from 12/27/16 to 5/24/17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order EC080967 has been initiated to update spraymation testing activities to include verification of placing the assembly machine back in service upon completion. This change should prevent missing components from being fully packaged. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of missing components/incorrect assembly. Cause of incorrect setup is covered. | 0701697203 | 0 | Y | United States | Closed | The customer reports one bag of syringes lot # 17A02563X had incomplete assembly of syringes - plungers were missing | AQUABILITI | 0012728353 | 2017-03-03 | 2017-03-23 | 0 | 2017-02-27 | During use | 2017-05-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Brooks, Janet COMPLAINT SOURCE OTHER: Aquabiliti Quality Engineer. ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Customer requires acknowledgement and close-out letters. Mitchell Eichas - e-mail - meichas@aquabiliti.com Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Customer still has product wating for RGA CONTACT NAME AND PHONE: Mitchell Eichas, phone: 615-833-2633 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event One bag of syringes lot # 17A02563X had incomplete assembly of syringes - plungers were missing. Was the customer notified the product should be returned? Yes SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet | Shanika Hodge | Alfredo Chavez | 42879 | 42879 | Product Analysis: One box/bag of 800 samples was returned for evaluation. The original product identification/shipping label was covered with a white label and thus lot number was unable to be confirmed. The box was intact and shipping damage was not observed. The package was opened and all 800 samples were visually inspected. Of 800 samples returned for evaluation, 159 were found to be missing a plunger rod and rubber tip (Incomplete assembled syringes). In the return bag sent by Aquabiliti the missing plungers could be observed as reported. | Complete | #N/A
2017-03-08 | 0701704048-10 | 8980 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8980 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: March 08, 2016 to April 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and 3 other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids packed prior to sealing a case. Also it unable to determine if the all missing lids cases are from the same lot number or different as lot number information is not provided. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701704048 | 0 | Y | United States | Closed | The customer reports ordered and received 7 cases of item #8980. When they opened the package there were 70 containers but zero lids. | 0 | 0 | 2017-03-08 | 2017-03-29 | 0 | 2017-03-08 | During use | 2017-06-06 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Trovato, Christopher COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Henry Schein ordered and received 7 cases of item #8980. When they opened the package there were 70 containers but zero lids. They would like the missing 70 lids to be sent out. Below are the details for Product - Model No:8980, Description:8GAL CONTAINER NEEDLE DISP X10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8GAL CONTAINER NEEDLE DISP X10 Model Number: 8980 Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No FACILITY NAME: Henry Schein FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 41 Weaver Road FACILITY CITY: Denver FACILITY STATE CODE: CO FACILITY POSTAL CODE: 17517 FACILITY DAY PHONE: 717)335-7230 ext4066 FACILITY EMAIL: raelin.schadler@henryschein.com SALESREP USER ID: trovac1 SALESREP EMP ID: 229312 SALESREP NAME: Trovato, Christopher PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Trovato, Christopher ADDITIONAL COMMENTS ABOUT THE EVENT: Please ship replacement lids to Raelin Schadler. | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-08 | 0701704056-10 | 8980 | 0174648 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL CONTAINER NEEDLE DISP X10 | MITG - Sharps Disposal | 002927 | 8980 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8980 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 0174648 (01/02/2017 to 01/12/2017) and no internal reject found during product ion. Trending Period: March 07, 2016 to April 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and 3 other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids packed prior to sealing a case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701704056 | 0 | Y | United States | Closed | The customer reports was shipped 30 sharps containers without lids. | HENRY SCHEIN ATHLETIC | 0012959921 | 2017-03-07 | 2017-03-28 | 0 | 2017-03-07 | During use | 2017-06-06 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Trovato, Christopher COMPLAINT SOURCE OTHER: Henry Schein Distributor Rep ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: The customer was shipped 30 sharps containers without lids. The customer is requesting replacement lids be sent. Below are the details for Product - Model No:8980, Serial/Lot No:0174648, Description:SharpSafety 8 Gallon Large Volume container w hiniged lkid IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for SharpSafety 8 Gallon Large Volume container w hiniged lkid Model Number: 8980 Serial or Lot Number: 0174648 ------------------------------------------------------------------------ ------------------------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No FACILITY NAME: Henry Schein FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 2556 Vista Blvd. #104 FACILITY CITY: Sparks FACILITY STATE CODE: NV FACILITY POSTAL CODE: 89434 FACILITY DAY PHONE: 717-335-7230 ext. 4066 FACILITY EMAIL: raelin.schadler@henryschein.com SALESREP USER ID: trovac1 SALESREP EMP ID: 229312 SALESREP NAME: Trovato, Christopher PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Trovato, Christopher ADDITIONAL COMMENTS ABOUT THE EVENT: Original PO #14287142 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-08 | 0701704547-10 | 31300919 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | SAG4838-TY 5 QT TURTUOUS PATH | MITG - Sharps Disposal | 002927 | 31300919 |  |  |  | RFR | PACKAGING | AFC | KC:PACKAGED WRONG PRODUCT | The Product 31300919 evaluated for the package wrong product. The reported condition cannot be observed due to no sample available for evaluation. DHR Review was performed and meet acceptance criteria. DHR Review: Device History Records data reviewed for 174541 (01/19/17 - 01/20/17) and no internal reject found during production for the same issue. Trending Period: Feb 28, 2016 to May 03, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause is label printer “print queue” may not be canceled before printing new label which may cause to print previous label from printer memory prior to print new labels. The human error and failure to check Overshipper label and product label to match before sealing a case. Formal Investigation: Refer to CAPA 2017-012626 for more details for corrective actions. Conclusion: This issue has been determined to be isolated event for 31300919 product. The product must be used as per product label identification. | 0701704547 | 0 | Y | Canada | Closed | The customer states the boxes show item 31300919 on the label but contain 8507SA inside the box. | Mills Memorial Hospital | 0001096246 | 2017-03-03 | 2017-03-22 | 0 | 2017-02-28 | Prior to Use | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Potential for Injury | 0 | Jacqueline St Pierre | Gatha Kothari | 42942 | 42942 | The sample photograph was reviewed. The Overshipper box had product code 31300919. The product inside was 8507SA. The reported condition was observed. Also lot number 174541 was obtained from the photo for DHR review. | Complete | #N/A
2017-03-08 | 0701704950-10 | 8881512852 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | LUBRICANT ISSUE | AFC | KC:EXCESSIVE LUBRICANT | Incident Description: The customer states when examining the syringe they found a greasy/slimy substance in the form of a dot on the rubber plunger. Product Analysis: Zero samples were returned for evaluation. A picture was available as seen below/attached in GCH. Based on the picture and part number provided, the syringe is not in its original primary packaging thus the tamper evident seal must have been broken. A slimy/greasy substance on the rubber plunger is observed in the form of a drop/dot. DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512852 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. The volume of silicone delivered is monitored weekly. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. The slimy/greasy substance reported is most probabilistic to be this silicone. Excess silicone is considered as a Minor defect in accordance with established quality and inspection standards. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size is 360 cases (172,800 pieces) with production occurring on production line 23. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: -Spraymation head levels were set incorrectly (Not significant – parameters were verified at each start up) -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective) Material Causes: -Use of incorrect silicone (Not significant – incoming inspection processes are in place and being followed) Method Causes: -Insufficient procedure for RAM cleaning (Significant – Cleaning/Wiping down of spraymation heads is not included daily cleaning activities. The heads have a potential of building up silicone during operation, with the potential of small drops falling into assembled syringes). -Insufficient preventative maintenance for Spraymation heads (Not Significant – A Preventative Maintenance plan is in place) -Process not validated (Not significant – spraymation process was confirmed to be validated) Measurement Causes: -QA inspection for spraymation testing not performed (Not significant – the inspection was confirmed to be performed as required) -Process parameters at start up were not verified (Not significant – parameters were verified) Machine Causes: -Spraymation heads not functioning properly/malfunction (Not Significant – The heads are confirmed to be on a regular rebuild schedule/PM) -Silicone delivery system not operating properly (Not significant – operation is confirmed via Syringe actuation testing throughout the lots) Environment Causes: -Environmental controls were out of validated operating ranges (Not significant – environmental controls are observed and documented once per shift) Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Excessive Lubricate in Syringe) and product family was performed. 3 other complaints of this nature for this product family were received in the trending period from 5.24.17 – 03.08.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.24.17 – 02.8.17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC023506 has been initiated to update daily cleaning activities, requiring cleaning/wiping down of spraymation heads. This cleaning will prevent build-up of silicone during manufacturing, decreasing the likelihood of silicone from falling into assembled syringes Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of excess silicone. Cause is due to normal manufacturing. Current process controls references Cleanliness GS20002641, which is the procedure which was updated. | 0701704950 | 0 | Y | Canada | Closed | The customer states when examining the syringe they found a greasy/slimy substance in the form of a dot on the rubber plunger. | YORKTON REGIONAL HEALTH CENTRE | 0012629903 | 2017-03-06 | 2017-03-23 | 0 | 2017-03-06 | Prior to Use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury - Potential for Injury | FIRST NOTIFIED NAME: Sparks, James richard ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Upon examination of the syringe greasy/slimy substance in the form of dot on the rubber plunger. Concerned about what this substance is, safe for use on patient? Below are the details for Product - Model No:8881512852, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Shewchuck PHYSICIAN FIRST NAME: Phyllis PHYSICIAN CLINIC: Yorkton Regional Health Centre PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Yorkton PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Colleen Bourne Regional Manager of Materials Management Yorkton Regional Health Centre 270 Bradbrooke Dr. Yorkton, SK Canada S3N 2K6 SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard ADDITIONAL COMMENTS ABOUT THE EVENT: This product complaint was issued by Sunrise, after a more general product alert was issued by 3S Health over the greasy/slimy substance on some 12ml syringes. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes Removed from Brief Description: Sunrise Yorkton 12ml Syringe March 2017 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-09 | 0701705049-10 | 8881512852 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:DIFFICULT TO BREAK SEAL | Incident Description: The customer states “These 12cc syringes are much more difficult for nursing staff to open. Previous is would take a light pressured twist or “butting” the base of the container against a desktop to open. This newest batch of 12 cc syringe containers are much harder/more difficult to open. Staff have reported their hands are getting sore from having to twist with more force or requiring additional “hitting/blows” of the lid against a stable surface to open.” Product Analysis: Samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and investigation may be performed. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size is 360 cases (172,800 pieces) with production occurring on production line 23. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: A defect could not be determined at this time. Nothing was discovered at the time of inspection or during the production run that would lead to the formulation of a root cause as product would not be released unless acceptance criteria is met. Torque testing is performed throughout the lot to ensure opening specifications are met. The customer may be attempting only method 1A (twisting the cap) and 1B (pressing the cap above the heat stake) as shown in the instructions for use below and provided with the product. It is advised that if the cap does not remove from twisting or pressing the edge of the cap above the heat stake with the thumb, or on a hard surface to break the seal such as in 1C. When the seal is broken, the cap should twist or pull off with applying minimum pressure. It should be noted that removal of the cap will be difficult if pushing inwards on the edges of the cap while attempting to remove it. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Difficult To Open/Difficult to break seal) and product family was performed. Zero other complaints of this nature for this product family were received in the trending period from 5.24.17 – 03.08.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.24.17 – 02.8.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. Further, other zero related non-conformances have been identified. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of heat staking with the associated controls for monitoring. | 0701705049 | 0 | Y | Canada | Closed | The customer states the syringes are more difficult than usual to open. | Saskatoon Health Region | 0015829658 | 2017-03-06 | 2017-03-23 | 0 | 2017-01-09 | During use | 2017-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Sparks, James richard ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512852, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No PHYSICIAN LAST NAME: Baliski PHYSICIAN FIRST NAME: Ryan PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Saskatoon PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Cote Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 FACILITY NAME: St. Pauls Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: Emergency SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes Removed from Brief Description: SPH 12 cc Syringe Opening Difficult | Jacqueline St Pierre | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-10 | 0701709213-10 | UNK NS | UNK | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | NEEDLES + SYRINGES UNKNOWN | MITG - Needles & Syringes | 003208 | UNK NS |  |  |  | RFR | BURN(S) | AFC | KC:UNIDENTIFIED LIQUID | Note Creator: Jeremy Roberts Note Date: 8/17/17 Note Description: Addition of Investigation Summary. Investigation Summary: Incident Description: Customer reports: Condensation inside syringe casing... when opened, the liquid has burned the skin of 3 individuals. Two subsequent complaints have been entered; one for each individual case. PE’s: 0701932743, 0701932744 and 0701709213. All complaints shall be investigated as one complaint. Product Analysis: As of 8/15/17 samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and investigation may be performed. Based on the reported information and pictures provided the following was noted: 1. GCH attachment “Condensate”, shows a hard pack luer lock syringe that has been removed from its primary packaging and reassembled/closed. The reported condition of condensation is observed in the sleeve. The Cap is shown on top of the Sleeve with a Luer Lock Syringe in the users hand. a. The only hard pack luer lock syringe made in Crystal Lake is 8881512878. b. 8881512878 is a sterile, hard pack, single use product. 2. GCH attachments “Thumb” and “thumb 2” show two separate thumbs with the incident description indicating the liquid has burned the skin of 3 individuals. a. This is an indication that the sterile single use syringe was used without gloves. It is not clear from the information provided if the same syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size for part 8881512878 is 360 cases and production line 23 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a process walk while performing a Six M assessment to identify areas where liquid is introduced into the manufacturing process that could make its way inside the molded sleeve. The assessment evaluated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Significant - Controlled Manufacturing Cleanliness GS20002641 does not have a quiz associated with the training item. Inclusion of a quiz will increase training effectiveness. GS20002641 outlines pre-cleaning activities requiring parts be removed from the manufacturing equipment. In the event that product remains in the Aidlin during periodic cleaning activities, cleaning agents can make their way into the sleeve and appear as condensation.) Material Causes: -Cleaning agent (Not Significant - Vesta Syde SQ Disinfectant is used in the manufacturing environment for cleaning on a periodic basis. Processes are in place to halt Manufacturing and protect product during cleaning activities. Burns or irritation to skin may occur if not properly used.) - Silicone (Not Significant – DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512878 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. In the event of a malfunction, silicone could be introduced into the sleeve and appear as condensation. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Burns or irritation to skin is unlikely). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and is required to be used during manufacturing). Method Insufficient Cleaning Process (Significant – Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. In the event solution is not properly diluted skin burns or irritation may occur.) Machine Causes: N/A: The only location within the manufacturing process where liquid is introduced with the possibility of entering the sleeve is during assembly where silicone is added to the syringe barrel. Reference material section above. Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: In the event of a failure during the manufacturing process, the most probable root cause is insufficient mixing of Vesta Syde SQ Disinfectant and Operator Training to GS20002641 Controlled Manufacturing Cleanliness GS20002641. It is not clear from the information provided if the syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Unidentified Liquid) and product family was performed. No other incidences of this nature for this product family were received in the trending period from 8.15.17 – 03.02.16. Zero internal rejects of this nature for this product family were received in the trending period from 8.15.17 – 12.2.16. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. In accordance with Medical Safety Review GCH transaction ID 226601208, Severity for this incident is a level 1. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. ComplianceWire training item CRY-GS20002641 Version 206.0.1 went effective 8/15/17 to active the quiz and initiate retraining for all affected employees. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure particulate and extraneous matter. | 0701709213 | 0 | Y | United States | Closed | Customer reports: Condensation inside syringe casing... when opened, the liquid has burned the skin of 3 individuals. Two subsequent complaints have been entered; one for each individual case. | SUMMIT VETERINARY REFERRAL CTR | 0013150347 | 2017-03-02 | 2017-03-10 | 0 | 2017-03-01 | During use | 2017-12-05 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Non-Serious Illness Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | Yes
2017-03-16 | 0701719439-10 | 8881412012 | 16M04463X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881412012 CURVED TIP SYRINGE X50 | MITG - Needles & Syringes | 003208 | 8881412012 |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:MOLDING FLASH | Incident Description: Syringes Do Not Plunge Smoothly Product Analysis: As of 5/26/17, samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and analysis may be performed. DHR Review: A review of the device history record for lot 16M04463X was completed. 39 (19,500 pieces) cases were completed on 01/8/2017 using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M assessment that addressed: Man, Mat erial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed as a corrective and preventative action when appropriate. Man Causes: -Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records). -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective). Material Causes: - Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met). - Rubber Tips (Not significant – An analysis is performed on tips at point of manufacturing to ensure all characteristics meet performance requirements. Incoming inspection is performed upon delivery of the tips to ensure requirements are met). Method Causes: Insufficient Standard Operating Procedure (Not Significant: Procedures define steps required for manufacturing. Validation is performed following standard operating procedures to as confirmation). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Machine Causes: Molding machine tip flash (Significant – Molding tip flash may create a partially occluded luer tip making it difficult to “Plunge Smoothly”) Assembly lubrication insufficient (Not Significant – The assembly silicone process is validated and monitored). Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Based on the investigation, the most probable root cause is that molding tip flash created a partial occlusion making it difficult to “Plunge Smoothly”. Molding flash is created when mold alignment components malfunction (worn luer tip/bushing). Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Molding Flash) and product family was performed. Two other complaints of this nature for this product family were received in the trending period from 5.26.17 – 02.24.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.26.17 – 01.24.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. Further, other zero related non-conformances have been identified. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed the failure/failure effect of flash with the associated cause of machine malfunctioning. The pFMEA confirms associated detection controls are in place for monitoring. | 0701719439 | 0 | Y | United States | Closed | the customer states ITEM# 8881412012 LOT# 16M04463X 5BX. SYRINGES DO NOT PLUNGE SMOOTHLY PO# F1590088. Patient involvement not provided. Discovered during testing. | BENCO DENTAL SUPPLY CO | 0012865979 | 2017-02-24 | 2017-03-16 | 0 | 2017-02-24 | Prior to Use | 2017-06-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-18 | 0701723188-10 | 8940FP | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8940FP SHARP TROLLEY MOLDED W/FOOT PEDAL | MITG - Sharps Disposal | 002927 | 8940FP |  |  |  | RFR | DOOR FAILS TO FUNCTION | AFC | KC:NON-FUNTIONAL CART | The Product 8940FP evaluated for the non-functional cart. The reported condition was observed during sample evaluation. DHR Review cannot be as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: March 16, 2016 to May 18, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the cart open-close mechanism for lid function. Root Cause: A possible root cause may be cart being damaged during shipping and handling causing cart sidewall warped inwards which causes resistance to lid function for cart. Also variation in the manufacturing process due to manual assembly of the cart. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as non-functional cart should not be used and replaced. | 0701723188 | 0 | Y | Canada | Closed | Per video: Lid fails to open automatically, needs to be done manually. | DELTA HOSPITAL | 0012614880 | 2017-03-16 | 2017-04-19 | 0 | 2017-03-16 | During use | 2017-11-04 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: ZS1 ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8940FP, Description:8940FP SHARP TROLLEY MOLDED W/FOOT PEDAL IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8940FP SHARP TROLLEY MOLDED W/FOOT PEDAL Model Number: 8940FP Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: demmy PHYSICIAN FIRST NAME: Jaquelynne PHYSICIAN CLINIC: Delta Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Delta PHYSICIAN STATE CODE: BC For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Dahlbeck, Michelle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZFOLCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH | Rose Matthews | Gatha Kothari | 42858 | 42873 | The sample was received and reported condition was observed. Also video was evaluated for sample evaluation. The lid open-close mechanism was working as foot pedal release. From the sample evaluation it seems like cart sidewall body was warped in which may cause external pressure on the container during use for lid open-close function. | Complete | Yes
2017-03-21 | 0701729570-10 | 31142222 | UNK | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 2 GAL CONTAINER; RED | MITG - Sharps Disposal | 002927 | 31142222 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31142222 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: March 21, 2016 to April 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and 4 other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids packed prior to sealing a case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701729570 | 0 | Y | United States | Closed | The customer states that they received a case of 31142222 but not lids were included. Requesting lids or RMA for return. Email sent to sales rep asking for account number, customer name of facility and email to facilitate RMA for product or SPO for lids. | 0 | 0 | 2016-11-14 | 2017-03-21 | 0 | 0 | Prior to Use | 2017-06-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Potential for Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-24 | 0701738878-10 | 1522SA | 00173931 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | PACKAGING | AFC | KC:CONTAMINATION | The Product 1522SA was evaluated for contamination. The reported condition was observed on the photos provided. DHR Review was performed and met the specification. DHR Review: Device History Record reviewed for the lot number 173931 (12/05/16 to 12/12/16). There was no internal reject found during production. Trending Period: Feb 14, 2016 – April 04, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection rejects and no other complaint for the product family. Root Cause: This issue has been determined to be an isolated event with minimal risk as contaminated product should not be used. Before use container/lid needs to be inspected for any abnormality. The potential root cause may be failure to inspect product for contamination prior to seal the finished good case. Also one of the photo shown the non-Medtronic Overshipper case which may be directly related to repackaging of the product outside of Medtronic which may resulted in to the damaged product and also contamination of product during handling due to human error. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as contaminated product should not be used. The product should be inspected prior to use for any abnormality. | 0701738878 | 0 | Y | United States | Closed | The customer states that: TWO OF THE SIX BOXES WERE PACKAGED DIFFERENTLY THAN IN THE PAST. INSIDE ONE OF THE ODDLY SHAPED BOXES, MOST OF THE SHARPS HAD STICKY RESIDUE ALL OVER THEM. CUSTOMER NEEDS A (COMPLAINT) RESPONSE. -###-From PE-701939124-The customer states that two of the six boxes were packaged differently than past. inside one of the oddly shaped boxes, most of the sharps had sticky residue all over them.-###-From SR-602067136-The customer states two of the six boxes were packaged differently than past. Inside one of the oddly shaped boxes, most of the sharps had sticky residue all over them. | Thermo Fisher Scientific | 0015890421 | 2017-02-14 | 2017-03-24 | 0 | 0 | Prior to Use | 2017-09-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-03-30 | 0701752403-10 | 065 | 00173917 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 065 1 QUART PHLEBOTOMY CONTAINER | MITG - Sharps Disposal | 002927 | 065 |  |  |  | RFR | LOW PUNCTURE FORCE | AFC | KC:NEEDLE PUNCTURED CONTAINER | The Product 065 evaluated for the needle punctured container. The reported condition was not observed as no photo or sample available for evaluation. DHR Review was performed and met the specification. DHR Review: Device History Records data reviewed for 173917 (10/27/16 to 10/30/16) and no internal reject found during production for the same. Trending Period: March 09, 2016 to May 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause is customer misuse. The container was filled above the fill line as per complaint description. There is a warning on the label not to over fill which was not followed during use. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event due to customer misuse. The container was filled above the fill line as per event description. There is no harm or injury reported for the complaint. | 0701752403 | 0 | Y | United States | Closed | The customer reports that the sharps container has been punctured. (Customer misuse)-###-From SR-601767471-The customer states; of hearing of a sharps container being punctured. it's clear from the warning on the label not to over fill it. the customer just wanted to make sure the manufacturer and you are aware of this in case it would isolated incident. | ROSE LINN CARE CTR | 0012523243 | 2017-03-09 | 2017-03-30 | 0 | 0 | During use | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | Yes
2017-04-04 | 0701761278-10 | 8612RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GA RCRA HAZARDOUS WASTE | MITG - Sharps Disposal | 002927 | 8612RC |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | KC:LID FIT ISSUE | The Product 8612RC evaluated for the lid fit issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: April 03, 2016 to Nov 30, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaint reported for the product family. Root Cause: The most probable root cause is that the R7413B tool has a larger rim height than the R7413A tool, which causes the lid locks to be difficult to secure around the rim. Formal Investigation: Refer to CAPA2017-012361 for more details for corrective and preventive action for R7413B container tool. CAPA is already implemented corrective action as planned. As for this complaint the lot number is unknown it cannot be determined when was the product manufactured. Conclusion: Refer to CAPA2017-012361 for more details for corrective and preventive action for R7413B container tool. | 0701761278 | 0 | Y | United States | Closed | The customer reports that the tops will not secure. This occurred during testing. | Memorial Hermann Surgery Ctr | 0001137031 | 2017-04-03 | 2017-04-04 | 0 | 2017-04-03 | Prior to Use | 2018-01-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Potential for Injury | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-04-04 | 0701761378-10 | 8982 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8982 evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record cannot be reviewed as lot number information is not provided. There was no internal reject found during production. Trending Period: March 29, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the container/lid style product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Also outside of Medtronic case may have been broken down to small quantity to ship to end user can cause missing absorbent pad issue. The absorbent pads are 4 packet of 5 each per packet (20 per case) to have one absorbent pad for each container as case quantity is 20 per case. There was no information in the complaint regarding quantity ordered by customer. The defective quantity listed as 1 each which means the customer did not receive full case of 20 eaches and case was broken down outside of Medtronic. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads. | 0701761378 | 0 | Y | United States | Closed | The Customer reports missing soaker pads | ST JOSEPHS MEDICAL SERVICE CTR | 0012834385 | 2017-03-29 | 2017-04-04 | 0 | 2017-03-29 | Unknown | 2017-06-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Justin Moss | 0 | 0 | 0 | 0 | Complete | #N/A
2017-04-05 | 0701763597-10 | 8881412012 | 16C04563X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881412012 CURVED TIP SYRINGE X50 | MITG - Needles & Syringes | 003208 | 8881412012 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:MOLDING FLASH | Incident Description: The diameter of the opening of the syringe is too small /narrow. Product Analysis: Two samples were returned for evaluation to analyze the reported condition. Visual inspection identified hard molding flash on the luer tip partially occluding the opening. The flash was then measured. Measurements confirmed flash was less than 1/32”, meeting specifications. The defect is not considered flash in the fluid path as it is hard flash making it an extension of the component part. DHR Review: A review of the device history record for lot 16C04563X and 16E03263X was completed. Each lot produced 240 cases (120,000 pieces) which were released on 04/13/2016 and 5/28/2016 respectively using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. DHR Review: A review of the device history record for lot 16C04563X and 16E03263X was completed. Each lot produced 240 cases (120,000 pieces) which were released on 04/13/2016 and 5/28/2016 respectively using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M assessment that addressed: Man, Mat erial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed as a corrective and preventative action when appropriate. Man Causes: -Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records). -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective). Material Causes: - Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met). - Rubber Tips (Not significant – An analysis is performed on tips at point of manufacturing to ensure all characteristics meet performance requirements. Incoming inspection is performed upon delivery of the tips to ensure requirements are met). Method Causes: Insufficient Standard Operating Procedure (Not Significant: Procedures define steps required for manufacturing. Validation is performed following standard operating procedures to as confirmation). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Machine Causes: Molding machine tip flash (Significant – Molding tip flash may create a partially occluded luer tip making a narrow opening) Assembly lubrication insufficient (Not Significant – The assembly silicone process is validated and monitored). Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Based on the investigation, the most probable root cause is that molding tip flash created a partial occlusion making it difficult to “narrow opening”. Molding flash is created when mold alignment components malfunction (worn luer tip/bushing). Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Molding Flash) and product family was performed. Three other complaints of this nature for this product family were received in the trending period from 5.26.17 – 04.04.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.26.17 – 03.04.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. Further, other zero related non-conformances have been identified. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. It has been noted that the associated molding tool is in process of being replaced. Replacement will occur in 2017. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed the failure/failure effect of flash with the associated cause of machine malfunctioning. The pFMEA confirms associated detection controls are in place for monitoring. | 0701763597 | 0 | Y | France | Closed | customer reports: the diameter of the opening of the syringe is too small /narrow | MEDLINE ASSEMBLY FRANCE SAS | 0001466636 | 2017-04-04 | 2017-04-26 | 0 | 2017-04-04 | Unknown | 2017-08-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Rose Matthews | Jeremy Roberts | 42885 | 42885 | Product Analysis: Two samples were returned for evaluation to analyze the reported condition. Visual inspection identified hard molding flash on the luer tip partially occluding the opening. The flash was then measured. Measurements confirmed flash was less than 1/32”, meeting specifications. The defect is not considered flash in the fluid path as it is hard flash making it an extension of the component part | Complete | #N/A
2017-04-05 | 0701763597-20 | 8881412012 | 16E03263X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881412012 CURVED TIP SYRINGE X50 | MITG - Needles & Syringes | 003208 | 8881412012 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:MOLDING FLASH | Incident Description: The diameter of the opening of the syringe is too small /narrow. Product Analysis: Two samples were returned for evaluation to analyze the reported condition. Visual inspection identified hard molding flash on the luer tip partially occluding the opening. The flash was then measured. Measurements confirmed flash was less than 1/32”, meeting specifications. The defect is not considered flash in the fluid path as it is hard flash making it an extension of the component part. DHR Review: A review of the device history record for lot 16C04563X and 16E03263X was completed. Each lot produced 240 cases (120,000 pieces) which were released on 04/13/2016 and 5/28/2016 respectively using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. DHR Review: A review of the device history record for lot 16C04563X and 16E03263X was completed. Each lot produced 240 cases (120,000 pieces) which were released on 04/13/2016 and 5/28/2016 respectively using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a 6M assessment that addressed: Man, Mat erial, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed as a corrective and preventative action when appropriate. Man Causes: -Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records). -Not following procedures correctly (Not significant – operators were trained to required procedures and training was found effective). Material Causes: - Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met). - Rubber Tips (Not significant – An analysis is performed on tips at point of manufacturing to ensure all characteristics meet performance requirements. Incoming inspection is performed upon delivery of the tips to ensure requirements are met). Method Causes: Insufficient Standard Operating Procedure (Not Significant: Procedures define steps required for manufacturing. Validation is performed following standard operating procedures to as confirmation). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Machine Causes: Molding machine tip flash (Significant – Molding tip flash may create a partially occluded luer tip making a narrow opening) Assembly lubrication insufficient (Not Significant – The assembly silicone process is validated and monitored). Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Based on the investigation, the most probable root cause is that molding tip flash created a partial occlusion making it difficult to “narrow opening”. Molding flash is created when mold alignment components malfunction (worn luer tip/bushing). Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Molding Flash) and product family was performed. Three other complaints of this nature for this product family were received in the trending period from 5.26.17 – 04.04.16. Zero internal rejects of this nature for this product family were received in the trending period from 5.26.17 – 03.04.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. Further, other zero related non-conformances have been identified. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. It has been noted that the associated molding tool is in process of being replaced. Replacement will occur in 2017. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed the failure/failure effect of flash with the associated cause of machine malfunctioning. The pFMEA confirms associated detection controls are in place for monitoring. | 0701763597 | 0 | Y | France | Closed | customer reports: the diameter of the opening of the syringe is too small /narrow | MEDLINE ASSEMBLY FRANCE SAS | 0001466636 | 2017-04-04 | 2017-04-26 | 0 | 2017-04-04 | Unknown | 2017-08-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Rose Matthews | Jeremy Roberts | 42885 | 42885 | Product Analysis: Two samples were returned for evaluation to analyze the reported condition. Visual inspection identified hard molding flash on the luer tip partially occluding the opening. The flash was then measured. Measurements confirmed flash was less than 1/32”, meeting specifications. The defect is not considered flash in the fluid path as it is hard flash making it an extension of the component part | Complete | #N/A
2017-04-06 | 0701768205-10 | 8900SA | 8900SA | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | LID ISSUE | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: April 04, 2016 to May 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaint reported for the product family which does not result in complaint trend yet. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8900SA is 100 eaches/case. The complaint is for 50 which shows the Medtronic case was broken down for further distribution in low unit of measure to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701768205 | 0 | Y | United States | Closed | The sharps containers were missing lids. 50 total | PATTERSON DENTAL SUPPLY | 0012865766 | 2017-04-06 | 2017-04-27 | 0 | 2017-04-05 | Prior to Use | 2017-06-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Cohen, Brian PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: ZS1 ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: The sharps containers were missing lids. 50 total. The customer is requesting to have them replaced. Below are the details for Product - Model No:8900SA, Serial/Lot No:8900SA, Description:8900SA 1QT RED CONTAINER PHLEBOTOMYX100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 Model Number: 8900SA Serial or Lot Number: 8900SA ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes SALESREP USER ID: cohenb4 SALESREP EMP ID: 328562 SALESREP NAME: Cohen, Brian PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Cohen, Brian | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-04-10 | 0701774984-10 | 31378089 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31378089 19 GAL SHARPS W/GASKET REDX5 | MITG - Sharps Disposal | 002927 | 31378089 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31378089 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: March 13, 2016 to May 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The complaint stated the product was repacked at third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids and/or containers should not be used. | 0701774984 | 0 | Y | United States | Closed | Lids were missing | Riverside Regional Medical Center | 0010763756 | 2017-03-13 | 2017-04-11 | 0 | 2017-03-13 | Servicing | 2017-06-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Hill, Lauren PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: 07 ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Sharps Container Lid Replacements Below are the details for Product - Model No:31378089, Description:31378089 19 GAL SHARPS W/GASKET REDX5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 31378089 19 GAL SHARPS W/GASKET REDX5 Model Number: 31378089 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No FACILITY NAME: Riverside Regional Med Ctr FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 500 J Clyde Morris Blvd FACILITY CITY: Newport News FACILITY STATE CODE: VA FACILITY POSTAL CODE: 23601 SALESREP NAME: Stanley, Teri PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Hill, Lauren ADDITIONAL COMMENTS ABOUT THE EVENT: Contacted from Owens and Minor Rep; Teri Stanley ¦ Customer Service Specialist ¦ Richmond DC 87 Owens & Minor ¦ 11900 N Lakeridge Pkwy, Ashland, VA 23005 t 757.534.6570 ¦ f 757.534.6572 ZFACILIT 00342554 Riverside Regional Med Ctr cannot find a BP in GCH | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-04-13 | 0701781165-10 | 8931 | 00173948 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8931 12GAL CONT CHEMO W/ WHI HNGD LIDX10 | MITG - Chemo Protection Safety | 003209 | 8931 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8931 evaluated for the missing lids. The reported condition was observed in photos and samples received. The total of 5 lids were missing in Non-Medtronic packaging. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 173948 (12/04/2016) and no internal reject found during production. Trending Period: April 12, 2016 to June 29, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8931 is 10 eaches/case. From sample product analysis it was discovered that Non-Medtronic packaging was used to ship product to end user. The repackaging of Medtronic case into Non-Medtronic packaging may be resulted in human error and failure to verify the required amount of lids for further distribution to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701781165 | 0 | Y | Hong Kong | Closed | The customer reported lids missing. | Hong Kong Baptist Hospital | 0001199709 | 2017-04-13 | 2017-05-02 | 0 | 2017-04-12 | Prior to Use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Yeung, Dave Event Context: Unknown, but will follow up for more information PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: ZP9 COMPLAINT SOURCE OTHER: Customer Warehouse's staff ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8931, Serial/Lot No:00173948, Description:8931 12GAL CONT CHEMO W/ WHI HNGD LIDX10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8931 12GAL CONT CHEMO W/ WHI HNGD LIDX10 Model Number: 8931 Serial or Lot Number: 00173948 ------------------------------------------------------------------------ ------------------------------------- When was the issue noticed? Before use/during setup Was the customer notified the product should be returned? Yes SALESREP USER ID: yeungw3 SALESREP EMP ID: 338769 SALESREP NAME: Yeung, Dave PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Yeung, Dave As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 338769 Yeung, Dave cannot find a BP in GCH ZINTCONT 338769 Yeung, Dave cannot find a BP in GCH ZINTCONT 338769 Yeung, Dave cannot find a BP in GCH ZFOLCONT 338769 Yeung, Dave cannot find a BP in GCH | Paul Deangelis | Gatha Kothari | 42915 | 42915 | The samples were received in Non-Medtronic Overshipper case. There are 10 containers and 5 lids in the samples. 5 lids are missing in the Non-Medtronic Overshipper case. | Complete | #N/A
2017-04-17 | 0701787028-10 | 8982 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8982 evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record cannot be reviewed as lot number information is not provided. There was no internal reject found during production. Trending Period: March 29, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and four other complaints reported for the container/lid style product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Also outside of Medtronic case may have been broken down to small quantity to ship to end user can cause missing absorbent pad issue. The absorbent pads are 4 packet of 5 each per packet (20 per case) to have one absorbent pad for each container as case quantity is 20 per case. There was no information in the complaint regarding quantity ordered by customer. The defective quantity listed as 2 each which means the customer did not receive full case of 20 eaches and case was broken down outside of Medtronic. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads. | 0701787028 | 0 | Y | United States | Closed | The customer reports missing soak pads | 0 | 0 | 2017-03-29 | 2017-04-17 | 0 | 0 | Unknown | 2017-06-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Justin Moss | 0 | 0 | 0 | 0 | Complete | #N/A
2017-04-24 | 0701802832-10 | 8982 | 00175232 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | LABELING ISSUE | AFC | KC:CLEAN NEEDLE STICK | The Product 8982 evaluated for the clean needle stick (used for chemo prep). The reported condition was observed in the photo. DHR Review was performed and met the specification. DHR Review: Device History Records data reviewed for 175232 (01/02/17 to 01/05/17) and no internal reject found during production. Trending Period: April 21, 2016 to May 22, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the similar product code. Root Cause: The potential root cause is customer misuse. The container was filled over ¾ of container per complaint description and improper handling of a filled sharps container by not holding on top for the disposal as needle punctured on the lower left side wall where it should not be held/lifted for further disposal. As there is no sample available for further evaluation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. The CDP-0011255 was initiated to add fill line on the 8982 - 2G YEL CHEMO CONTAINER label as corrective/preventative action. Conclusion: This issue has been determined to be an isolated event due to customer misuse. The container was above ¾ full and it was lifted from the bottom/sides not by using handle or top part of the container. | 0701802832 | 0 | Y | United States | Closed | The Customer states: that there was an injury by a needle from a container and said that the container was not labled like the red containers saying how much to fill them. to avoid injuries. The container contents were above the fill line. The needle protruding from the container was not visible when the staff member when to pick the container up for disposal, resulting in a needle stick to the hand. The protruding needle was not used on a patient, it was used for chemo prep. The staff member has blood work and treatment, but the specifics cannot/will not be shared. | OLIVE VIEW MEDICAL CTR | 0013184263 | 2017-04-21 | 2017-04-24 | 0 | 2017-04-21 | Servicing | 2017-07-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Serious Illness Injury | 0 | Daniel Tamul | Gatha Kothari | 42877 | 42877 | The photo was received in complaint. The needle was protruding on lower left side wall when container label is facing front. As per complaint event description container contents were above the ¾ of container. There is no physical sample available for further analysis. | Complete | Yes
2017-04-26 | 0701807545-10 | 31142222 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 2 GAL CONTAINER; RED | MITG - Sharps Disposal | 002927 | 31142222 |  |  |  | RFR | CONTAINER ISSUE | AFC | KC:DAMAGED PRODUCT | The Product 31142222 evaluated for the damaged product. The reported condition was not observed as no sample or photo available for evaluation. DHR Review was not performed due to unknown lot number. DHR Review: A review of the Device History Record cannot be conducted since a lot number was unknown. Trending Period: April 03, 2016 to May 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The potential root cause may be product (container/lids) may have been damaged during shipping and handling due to external stress/forces. As per the complaint no. of defective quantity is 6 eaches, which shows that case may had been broken down to small quantity by distribution center. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage as per IFU. If damaged found product should not be used and must be replaced. | 0701807545 | 0 | Y | United States | Closed | the customer reports product is damaged and unusable. | CINCINNATI SPORTS MEDICINE | 0013199403 | 2017-04-03 | 2017-04-26 | 0 | 2017-04-03 | Unknown | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Stacey Olsen | 0 | 0 | 0 | 0 | Complete | #N/A
2017-04-27 | 0701812385-10 | 8607RC | 00172637 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8 GAL RCRA HINGE TOP | MITG - Sharps Disposal | 002927 | 8607RC |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8607RC evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 172637 (10/17/16 to 10/19/16). There was no internal reject found during production. Trending Period: April 26, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the container/lid style product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads and other label kit items packed prior to sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads and other kit items should not be used. | 0701812385 | 0 | Y | United States | Closed | The Customer stated missing plastic bags that contain Soaking Pad, Red Sticker and White Stickers, signifying Hazardous Materials for each container. | Medline Industries, Inc. | 0011846340 | 2017-04-26 | 2017-04-27 | 0 | 0 | During use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-01 | 0701816028-10 | 8900SA | 00174639 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:LABEL ISSUE | The Product 8900SA evaluated for the label issue. The reported condition was observed on the physical sample which does not affect form, fit and functionality. DHR Review performed and met acceptance criteria. DHR Review: Device History Record reviewed for the lot number 174639 (02/07/17 to 02/17/17) and 175501 (03/08/17 to 03/14/17). There was no internal reject found during production. Trending Period: May 09, 2016 – May 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The 8900SA containers are molded, labeled and stacked during production process. There may be potentially label was not properly stick to the container and while stacking and packing one of the label corner slightly peeled off. The potential root cause is human error and failure to inspect label peel of before packaging in to the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as all label information is well legible and except one corner label was properly adhered to the container, therefore does not pose any harm to the end user. | 0701816028 | 0 | Y | Japan | Closed | The customer reports labeling was found peeled off prior to use. | Niti-on Co., Ltd | 0015877231 | 2017-05-01 | 2017-05-01 | 0 | 2017-04-30 | Prior to Use | 2017-06-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | Gatha Kothari | 42879 | 42879 | There are two samples received for evaluation. The lot no. from samples are 174639 and 175501 one each for 8900SA. The bottom right corner slightly peeled off on 174639 lot number sample. The bottom left corner slightly peeled off on 175501 lot number sample. In both samples product code, lot number and all other information are well legible. The reported condition observed but it will not affect form, fit and functionality of the part. The label was properly adhered to the container except slightly peeled on one corner. | Complete | #N/A
2017-05-02 | 0701819699-10 | 8850 | 00174333 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8GAL.PHARMASAFETY WASTE CNT.10 | Unknown |  | 8850 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8850 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174333 (01/19/2017 to 01/21/2017) and no internal reject found during production. Trending Period: March 30, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8850 is 10 eaches/case. The complaint is for 4 eaches which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701819699 | 0 | Y | United States | Closed | The Customer states missing lids | Medline Industries, Inc. | 0011846340 | 2017-03-30 | 2017-05-02 | 0 | 0 | Unknown | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-03 | 0701821824-10 | 8938 | 00176259 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING COMPONENTS | The Product 8938 evaluated for the missing component (lid component). The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176259 (03/20/17 to 03/22/17) and no internal reject found during production. Trending Period: May 02, 2016 to May 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: There are three components of Lid assembly (Lid, Door and Rail). The potential root cause may be one of the component was not assembled in the manual assembly process. As there is no sample or photo available to evaluate further. This issue has been determined to be an isolated event with minimal risk. Before use container/lid needs to be inspected per IFU. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing component (lid component) product should not be used. | 0701821824 | 0 | Y | United States | Closed | The Customer states sharps cont slide/top is missing upon opening box | HENRY SCHEIN INC | 0013191914 | 2017-05-02 | 2017-05-03 | 0 | 0 | Unknown | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-03 | 0701822140-10 | V8985 | 00175668 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | V8985 NOVAPLUS 8GA HINGED CHEMO-YEL | MITG - Sharps Disposal | 002927 | V8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product V8985 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175668 (03/23/2017 to 03/25/2017) and no internal reject found during production. Trending Period: May 02, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for V8985 is 10 eaches/case. The complaint is for 1 each which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701822140 | 0 | Y | United States | Closed | The Customer states : missing lids | JOHN FITZGIBBON MEMORIAL HOSP | 0013178518 | 2017-05-02 | 2017-05-03 | 0 | 0 | Unknown | 2017-06-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-03 | 0701822940-10 | 8506Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8506Y YEL 5 QRT CONT S STAR C B DR X20 | MITG - Sharps Disposal | 002927 | 8506Y |  |  |  | RFR | LID ISSUE - FAILURE TO DROP SHARP | AFC | KC:MALFUNCTIONING LID | The Product 8506Y evaluated for the failure to drop sharp-malfunction lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: April 09, 2016 to May 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: As per the communication from the sales rep the container was not full above the fill line. The Winged Infusion Sets with flexible tubing and large syringes are being disposed. The tubing get stuck in the back of the container lid and the lid will no longer open once this happens. Also large syringes do not fit easily in the small lid opening and are hard to drop and cause the lid to get stuck when dropping. As from the information provided it seems like; it may require different style of lid-container should be used in this type of setting. The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. It may be different style lid/container opening unit should be used for customer’s specific end use application. | 0701822940 | 0 | Y | Canada | Closed | customer reports: Sharps container is jamming so staff are being exposed to previous used sharps when attempting to dispose of other sharps. | Langley Memorial Hospital | 0001097292 | 2017-05-01 | 2017-05-04 | 0 | 2017-04-09 | During use | 2017-08-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: ZS1 ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: "Sharps container is jamming so staff potentially are being exposed to previous used sharps when attempting to dispose of other sharps." - I am meeting with Raemone to discuss what is happening and if this is actually product quality issue or if the containers are just to full and are "jamming" because of this. Below are the details for Product - Model No:8506Y, Description:8506Y YEL 5 QRT CONT S STAR C B DR X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8506Y YEL 5 QRT CONT S STAR C B DR X20 Model Number: 8506Y Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? During use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Friesen PHYSICIAN FIRST NAME: Raemone PHYSICIAN CLINIC: Langley Memorial Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Langley PHYSICIAN STATE CODE: BC PHYSICIAN DAY PHONE: 604-514-6074 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Dahlbeck, Michelle ADDITIONAL COMMENTS ABOUT THE EVENT: I am meeting with Raemone on Friday to confirm whether this is an actual product quality issue or a user issue. Will send follow up information after meeting As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZFOLCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-05-05 | 0701827739-10 | 8938 | 00176259 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8938 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176259 (03/20/2017 to 03/22/2017) and no internal reject found during production. Trending Period: May 04, 2016 to May 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8938 is 5 eaches/case. The complaint is for 6 which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701827739 | 0 | Y | United States | Closed | The Customer stated received Containers without Lids. | CARDINAL HEALTH | 0012924700 | 2017-05-04 | 2017-05-05 | 0 | 2017-04-27 | Prior to Use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-08 | 0701830973-10 | 8931 | 00175358 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8931 12GAL CONT CHEMO W/ WHI HNGD LIDX10 | MITG - Chemo Protection Safety | 003209 | 8931 |  |  |  | RFR | LID ISSUE | AFC | KC:MISSING LIDS | The Product 8931 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175358 (02/07/2017 to 02/08/2017) and no internal reject found during production. Trending Period: April 21, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8931 is 10 eaches/case. The complaint is for 1 each which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701830973 | 0 | Y | United States | Closed | the customer reports missing lids. | 0 | 0 | 2017-04-21 | 2017-05-08 | 0 | 2017-04-21 | Unknown | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Stacey Olsen | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-08 | 0701831550-10 | 8881112599 | 17C03663X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112599 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NEEDLE ISSUE | AFC | KC:DAMAGED COMPONENTS | Incident Description: The customer reported defective dented syringes. Syringes have 1-2 dents stopping the line and machines have to be reset. 1 out of every 25 syringes found with defect. Product Analysis: Three samples have been returned for evaluation. Upon examination of the syringe subassembly, the syringes were observed to be dented as reported. DHR Review: A review of the device history record for lot 17C03663X was completed. 624 cases were completed on 03/20/2017 using production line 1. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Assembly System Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Material Causes: Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met. Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Method Causes: Insufficient Standard Operating Procedure (SOP) for setting/troubleshooting the Aidlin. (Not Significant – Method is sufficient based on trending results). Machine Causes: Conveyer belt jamming (Not significant – Part(s) jammed in conveyer belt are not likely to exhibit defects similar to the one being investigated. Part(s) would likely trigger high plungers by sensors at the RAM and not make it through the automated assembly system.) Feeding rail of the Aidlin (Significant – Root Cause, Burrs on feeding rail made contact with the parts.) Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Based on the product and root cause analysis, the root cause of the defect due to damaged feeding rail on Aidlin AM-240 contacted the syringe as it rotated in the star-wheel. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Damage Functional Less Than Desired) and product family was performed. Zero other complaints of this nature for this product family were received in the trending period from 05.08.16 – 06.09.17. Zero internal rejects of this nature for this product family were received in the trending period from 10.30.16 – 06.09.17. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. The below improvement actions were previously completed on 04/13/2017. Furthermore, zero other related non-conformances have been identified during trending. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: MAXIMO Work Order CL-2521330 was created on 04/12/2017 and completed on 04/13/2017. This Work Order corrected the feeding rail which was determined to be the root cause of the defect under investigation. Preventative maintenance for AM-240-Q was implemented on 04/09/2015 and is sufficient based on trending results. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure and root cause for “Damage Functional Less Than Desired” is reflected. | 0701831550 | 0 | Y | United States | Closed | The customer reported defective dented syringes. The syringes have 1-2 dents stopping the line and the machines have to be reset. 1 out of every 25 syringes found with this defect. | AQUABILITI | 0012728353 | 2017-05-08 | 2017-05-17 | 0 | 2017-05-08 | During use | 2017-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Brooks, Janet PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Regulatory/Competent Authority ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Syringes have 1-2 dents stopping the line and machines have to be reset. 1 out of every 25 syringes found with defect. Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Product still at Aquabiliti - QA Crystal Lake will have to approve return. CONTACT NAME AND PHONE: Mitchell Eichas, Phone: 615-833-2633 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event Lot# 17C03663X - Defective dented syringes. Was the customer notified the product should be returned? Yes SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer sent picture and I have requested samples. Customer requires written close-out letter. | Janice Nevius | Juan Diaz | 42873 | 42898 | Three samples have been returned for evaluation. Upon examination of the syringe subassembly, the syringes were observed to be dented as reported. | Complete | #N/A
2017-05-09 | 0701834217-10 | 8507SA | UKNOWN | COV LP LLC Crystal Lake DC | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: May 08, 2016 to May 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 8507SA is 20 eaches/case. The complaint is for 1 each which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701834217 | 0 | Y | United States | Closed | The customer states missing lid . | Medline Industries, Inc. | 0011846340 | 2017-05-08 | 2017-05-09 | 0 | 0 | During use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-11 | 0701838060-10 | 8939 | 00175302 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | CONTAINER ISSUE | AFC | KC:COSMETIC ISSUE | The Product 8939 evaluated for the cosmetic issue – stress mark. The reported condition was observed on the physical sample but it is well within cosmetic specification criteria. DHR Review performed and met acceptance criteria. DHR Review: Device History Record reviewed for the lot number 175302 (02/11/17 to 02/14/17). There was no internal reject found during production. Trending Period: May 09, 2016 – May 24, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family/lid type. Root Cause: The sample evaluation was performed on returned samples. As per GS10000913 stress mark is allowed on surface B less than 2” in length or up to 1/10 of part length whichever is less and less than 2 marks in total as reaction plan. If there are more than 5 stress marks greater than 2” in length or greater than 1/10 of the length of the part than it is a reject. The samples received have only one stress mark per container well within reaction criteria. The reported condition stated by the customer was not observed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as visual-cosmetics are well within acceptance criteria, therefore does not pose any harm to the end user. | 0701838060 | 0 | Y | Japan | Closed | The customer reports scratch mark was found on the product prior to use. | Niti-on Co., Ltd | 0015877231 | 2017-05-10 | 2017-05-11 | 0 | 2017-05-09 | Prior to Use | 2017-06-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | Gatha Kothari | 42879 | 42879 | There are total 5 samples (1 case) were submitted with this complaint. As per GS10000913 stress mark is allowed on surface B less than 2” in length or up to 1/10 of part length whichever is less and less than 2 marks in total as reaction plan. If there are more than 5 stress marks greater than 2” in length or greater than 1/10 of the length of the part than it is reject. The sample received has one stress mark in total per container well within reaction criteria. The reported condition stated by the customer was not observed. | Complete | #N/A
2017-05-11 | 0701840106-10 | 8938 | 1000176259 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | LID ISSUE | AFC | KC:NO FAILURE | The Product 8938 evaluated for “no failure” - the complaint’s event description which is very limited to perform any investigation. The reported condition cannot be observed as no photo or physical sample or more information available for evaluation. DHR Review performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176259 (03/20/17 to 03/22/17) and no internal reject found during production. Trending Period: May 02, 2016 to May 30, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family for “sticking”. Trending cannot be performed due to lack of information in the complaint event description. Root Cause: The event description noted in this complaint, has been confirmed by Medtronic PMV complaint handling to be the only information that was available from the complainant. As there was no available clarification of the complaint event, the Crystal Lake manufacturing site was limited in the investigation. As there is no sample/photo or more information available for further evaluation of the event. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: Based on the above information and the minimal information present in the complaint, record, there is no confirmation of any defect events associated with ‘sticking” or “not sticking” for this product. | 0701840106 | 0 | Y | United States | Closed | THE CUSTOMER STATES NOT STICKING | HENRY SCHEIN INC VETERINARY | 0013106500 | 2017-05-04 | 2017-05-11 | 0 | 2017-05-02 | Servicing | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Unknown | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-11 | 0701840148-10 | 8980Y | 168768 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980Y YEL 8 GAL SHARPS CONT HINGED LID | MITG - Sharps Disposal | 002927 | 8980Y |  |  |  | RFR | PACKAGING | AFC | KC:MISSING COMPONENTS | The Product 8980Y evaluated for the missing component. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 168768 (12/01/15 to 12/04/15) and no internal reject found during production. Trending Period: May 08, 2016 to May 31, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of components presents before sealing the case. There is no information about which components are missing in the complaint. Also there are no samples or photos available for evaluation. The Medtronic case quantity for 8980Y is 10 eaches/case. The complaint quantity is 1 each. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. This issue has been determined to be an isolated event with minimal risk. Before use product needs to be inspected per IFU. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing component product should not be used. | 0701840148 | 0 | Y | Chile | Closed | The customer reports packaging has incorrect quantity and components are loose. | PONTIFICIA UNIVERSIDAD CATOLICA DE CHILE | 0013036127 | 2017-05-11 | 2017-05-12 | 0 | 2017-05-08 | Prior to Use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | No Product found for LotSerial# 168768 No Product found for LotSerial# 168768 FIRST NOTIFIED NAME: Medina estrada, Juan carlos PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No REPORTED TO AUTHORITY: No Below are the details for Product - Serial/Lot No:168768, Description:KITCHCABEZAYCUELLOPP IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for KITCHCABEZAYCUELLOPP Model Number: Not specified Serial or Lot Number: 168768 ------------------------------------------------------------------------ ------------------------ Please select all that apply Packaging issue Please select all that apply Package contains incorrect quantity Please specify sin insumo , 1 unidad código E1455 Was the customer notified the product should be returned? Yes Below are the details for Product - Serial/Lot No:168768, Description:KITCHCABEZAYCUELLOPP IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for KITCHCABEZAYCUELLOPP Model Number: Not specified Serial or Lot Number: 168768 ------------------------------------------------------------------------ ------------------------ Please select all that apply Component issue Please select all that apply Component loose Please specify sin unidad código E1455 Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Becker PHYSICIAN FIRST NAME: Javiera PHYSICIAN CLINIC: Hosp. Univ. Catolica PHYSICIAN COUNTRY CODE: CHL PHYSICIAN ADDRESS1: Marcoleta 367 PHYSICIAN CITY: Santiago PHYSICIAN PROVINCE: RM PHYSICIAN DAY PHONE: 223543526 PHYSICIAN EMAIL: jbecker@med.puc.cl For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: estraj13 SALESREP EMP ID: 289042 SALESREP NAME: Medina estrada, Juan Carlos PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Medina estrada, Juan carlos As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-17 | 0701849887-10 | 89651 | 173203 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 89651 2GAL RED CONT H DROP CLR LIDX20 | MITG - Sharps Disposal | 002927 | 89651 |  |  |  | RFR | PACKAGING | AFC | KC:CRACKED LIDS | The Product 89651 evaluated for cracked lid. The reported condition cannot be observed as no photo or physical sample available for evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 173203 (08/10/16 to 08/13/16) and no internal reject found during production. Trending Period: May 16, 2016 to June 02, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: This issue has been determined to be an isolated event with minimal risk as cracked lids should not be used. Before use container/lid needs to be inspected for any damage per IFU. The potential root cause has been determined to be damage during shipping and handling due to external stress/forces. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as cracked lid product should not be used. The issue was identified prior to use and there is no harm or injury reported to patient safety. | 0701849887 | 0 | Y | United States | Closed | The customer reported cracked lids | 0 | 0 | 2017-05-16 | 2017-05-19 | 0 | 2017-05-16 | Prior to Use | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | FIRST NOTIFIED NAME: Razzi, Elisabeth Event Context: Asked but unknown PATIENT MEDICAL HISTORY: Asked but unknown PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Unknown Asked but unknown Did the event lead to or extend patient hospitalization? Unknown Asked but unknown Was there any injury as a result of the event? No injury If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: Yes PRODUCT EVENT DETAILS OTHER: The customer received 4 cracked lids, please send replacement lids ( 4 eaches). Below are the details for Product - Model No:89651, Serial/Lot No:173203, Description:89651 2GAL RED CONT H DROP CLR LIDX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 89651 2GAL RED CONT H DROP CLR LIDX20 Model Number: 89651 Serial or Lot Number: 173203 ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth ADDITIONAL COMMENTS ABOUT THE EVENT: Please send replacement lids to: Unimed Corp 100 Elm Street – Unit 300D Warren, RI 02885 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-17 | 0701849941-10 | 8939 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | PACKAGING | AFC | KC:DAMAGED PRODUCT | The Product 8939 evaluated for the damaged product. The reported condition was not observed as no sample or photo available for eval uation. DHR Review was not performed due to unknown lot number. DHR Review: A review of the Device History Record cannot be conducted since a lot number was unknown. Trending Period: May 08, 2016 to May 31, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces. As there is no sample or photos available to perform further investigation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage as per IFU. If damaged found product should not be used and must be replaced. | 0701849941 | 0 | Y | United States | Closed | The customer states Mid state Medical Center received 1 CS of # 8939. Upon opening the outer case the receiving department noticed there was one container that was damaged and deemed unusable. The customer is looking to replace 1 EA of the 8939 container.” | Midstate Medical Center | 0015734705 | 2017-05-15 | 2017-05-17 | 0 | 2017-05-08 | Unknown | 2017-07-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-17 | 0701850302-10 | 85221R | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85221R 3GAL TRANS RED CONTAINER CLRTOP | MITG - Sharps Disposal | 002927 | 85221R |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:LID FIT ISSUE | The Product 85221R evaluated for lid fit issue. The reported condition cannot be observed as no photo or physical sample or more information available for evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: May 16, 2016 to June 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family for the same. Root Cause: The root cause cannot be determined as no sample or photos available for product analysis. Also DHR review cannot be performed as lot no. is unknown. The potential cause may be user error to follow guideline for lid fit per IFU (Instruction for Use). Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: Based on the above information and the minimal information present in the complaint and lack of sample/photo defect was not observed. | 0701850302 | 0 | Y | United States | Closed | The customer received lids with the 3 gallon sharp container, but the lids do not fit the containers. | MOUNTAINVIEW MEDICAL CTR | 0012873805 | 2017-05-16 | 2017-05-17 | 0 | 0 | Prior to Use | 2017-08-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-05-22 | 0701859607-10 | 31314886 | UKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31314886 3 GAL TRANS RED CONTAIN SHARPS | MITG - Sharps Disposal | 002927 | 31314886 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31314886 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: May 09, 2016 to June 02, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during repackaging at the third party distribution. The Medtronic case quantity for 31314886 is 12 eaches/case. The complaint is for 6 eaches which shows the Medtronic case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids and/or containers should not be used. | 0701859607 | 0 | Y | United States | Closed | The customer states missing lids. | Medline Industries, Inc. | 0011846340 | 2017-05-09 | 2017-05-22 | 0 | 0 | Unknown | 2017-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-23 | 0701861034-10 | 8881512852 | 16J06563X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | SHEATH/CAP ISSUE | AFC | KC:DIFFICULT TO OPEN PACKAGE | Incident Description: The customer reports caps not popping off as easily as normal, requiring a lot more force. Product Analysis: One sample was returned for evaluation. Primary packaging on the sample was already opened which did not allow for torque testing to be performed. The sample was visually inspected for defects. No defects were observed that could contribute to the reported condition. DHR Review: A review of the device history record for lot 16J06563X was completed. 360 cases (172,800 pieces) were produced and released on 11/2/16 using production line 23. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: A defect could not be determined at this time. Nothing was discovered at the time of inspection or during the production run that would lead to the formulation of a root cause as product would not be released unless acceptance criteria is met. Torque testing is performed throughout the lot to ensure opening specifications are met. The customer may be attempting only method 1A (twisting the cap) and 1B (pressing the cap above the heat stake) as shown in the instructions for use below and provided with the product. It is advised that if the cap does not remove from twisting or pressing the edge of the cap above the heat stake with the thumb, or on a hard surface to break the seal such as in 1C. When the seal is broken, the cap should twist or pull off with applying minimum pressure. It should be noted that removal of the cap will be difficult if pushing inwards on the edges of the cap while attempting to remove it. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to this failure mode (Difficult To Open/Difficult to break seal) and product family was performed. One other complaint of this nature for this product family were received in the trending period from 6.12.17 – 05.09.16. Zero internal rejects of this nature for this product family were received in the trending period from 6.12.17 – 04.9.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. Further, other zero related non-conformances have been identified. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed the failure/failure effect of heat staking with the associated controls for monitoring. | 0701861034 | 0 | Y | Canada | Closed | The customer reports caps not popping off as easily as normal, requiring a lot more force. | 0 | 0 | 2017-05-19 | 2017-05-24 | 0 | 2017-05-16 | During use | 2017-08-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Sparks, James richard PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Caps not popping off as easily as normal, requiring a lot more force. Below are the details for Product - Model No:8881512852, Serial/Lot No:16J06563X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 16J06563X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? Yes Provide details Minor injury to staff Was the customer notified the product should be returned? Yes For Physician with Siebel ID 0015935727: PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Côté RN ORN BSN Ph. 306-655-1654 Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 Kara.cote@saskatoonhealthregion.ca FACILITY NAME: St. Pauls Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: Hemodialysis CCRHC SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0015935727 Jennifer Larsen cannot find a BP in GCH | Paul Deangelis | Jeremy Roberts | 42943 | 42943 | One sample was returned for evaluation. Primary packaging on the sample was already opened which did not allow for torque testing to be performed. The sample was visually inspected for defects. No defects were observed that could contribute to the reported condition. | Complete | #N/A
2017-05-23 | 0701861125-10 | 8881512878 | 17C02363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | CONTAMINATION | AFC | KC:UNIDENTIFIED MATTER | Incident Description: The customer states: Brown substance on the plunger. The brown substance was on the outside of the plunger on the circular portion that is visible once the protective cap has been removed, not in the fluid path. Product Analysis: One physical sample received for evaluation. The sample exhibits a brown substance as described by the customer. The brown substance is predominantly on outer rim of the circular end of the plunger and scattered areas throughout the body of the plunger. DHR Review: A review of the device history record for lot 17C02363X has been completed. This product was manufactured March 17-19, 2017. Review of the device history record did not identify any manufacturing challenges which could lead to this alleged defect during the production of these lots. All acceptance criteria were met prior to final release. Review of these lot records show that all statistically based sampling visual, packaging and functional passed allowing the product to be released meeting the appropriate quality criteria. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Material Causes: Resin (Not significant – An analysis is performed on Resin at point of manufacturing to ensure all characteristics meet performance requirements. A certification is received upon delivery of the resin to ensure requirements are met) Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Method Insufficient Standard Work: Molding area inspection responsibilities for Syringe Focus Factory Manufacturing Associates. (Significant – Probable Root Cause, as per Standard Work, Manufacturing associates are not required to examine press for loose parts during start-up shots providing an opportunity for scrap to enter normal production). Machine Causes: Process Parameters (Not significant – Discoloration due to irregular process parameters is not representative of the defect under in vestigation) Environment Causes: N/A: Production is completed in a controlled manufacturing environment. The probable root cause of contamination is believed to have occurred during the molding process start up. The investigation team concluded the initial shots during molding process start-up can exhibit contamination defects resembling the contamination being investigated. As per the standard operating procedure PS20001913 Rev. 4, a minimum of 10 shots is required before inspection shot samples are pulled. All shots prior to the manufacturing of usable parts are run into a white usable scrap brute. The press is not inspected for any parts that may have bounced out of the typical process and avoided being properly discarded into a usable scrap brute. The loose parts remaining in the press can be reintroduced into the process later in the run. Trending Analysis: A review of customer complaints as well as internal rejects pertaining to the failure mode (Contamination) with related root cause and product family was performed. Zero complaints of this nature for this product family were received in the trending period from 03.01.16 – 07.07.17. Zero internal rejects of this nature for this product family were received in the trending period from 02.01.17 – 07.07.17. Formal Investigation: Based on the existing controls, the internal reject process and the complaint history review, no formal investigation is required at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventative Action: Modification will be made to DC00043215 Rev. 0. These updates will instruct manufacturing associates to ensure the press and conveyer are clear of all parts manufactured during the start-up process prior to manufacturing usable parts. Aforementioned updates will prevent contaminated start-up samples from being introduced into the normal process and prevent reoccurrences of this defect. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Value Stream – Molding was reviewed which confirmed the failure and root cause for Contamination/Specs (Non-Fluid Path) is reflected. | 0701861125 | 0 | Y | Canada | Closed | The customer reported brown substance on the plunger. The brown substance was on the outside of the plunger on the circular portion that is visible once the protective cap has been removed, not in the fluid path. | Saskatoon Health Region | 0008045589 | 2017-05-23 | 2017-05-24 | 0 | 2017-05-17 | During use | 2017-08-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Sparks, James richard PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512878, Serial/Lot No:17C02363X, Description:8881512878 12CC S-LK SYR ONLY O/P X480 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881512878 12CC S-LK SYR ONLY O/P X480 Model Number: 8881512878 Serial or Lot Number: 17C02363X ------------------------------------------------------------------------ ------------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Gabriel PHYSICIAN FIRST NAME: Aaron PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Saskatoon PHYSICIAN STATE CODE: SK For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Kara Côté RN ORN BSN Ph. 306-655-1654 Manager Clinical Equipment & Product Standardization Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 Kara.cote@saskatoonhealthregion.ca FACILITY NAME: Saskatoon City Hospital FACILITY COUNTRY CODE: CAN FACILITY CITY: Saskatoon FACILITY STATE CODE: SK As Reported Event - departmentName: 3100/3200/3300 SALESREP USER ID: sparkj2 SALESREP EMP ID: 252398 SALESREP NAME: Sparks, James Richard PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Sparks, James richard As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Paul Deangelis | Jeremy Roberts | 42923 | 42923 | One physical sample received for evaluation. The sample displayed a brown substance as described by the customer. The brown substance is predominantly on outer rim of the circular end of the plunger and scattered areas throughout the body of the plunger. | Complete | #N/A
2017-05-24 | 0701863530-10 | V8985 | UKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | V8985 NOVAPLUS 8GA HINGED CHEMO-YEL | MITG - Sharps Disposal | 002927 | V8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product V8985 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: May 01, 2016 to June 05, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The Medtronic case quantity for V8985 is 10 eaches/case. The complaint is for 1 each which shows that Medtronic case may be broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids and/or containers should not be used. | 0701863530 | 0 | Y | United States | Closed | The customer states : missing lids. | JOHN FITZGIBBON MEMORIAL HOSP | 0013178518 | 2017-05-01 | 2017-05-24 | 0 | 0 | During use | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-29 | 0701870766-10 | 1522SA | 175631 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | CONTAMINATION | AFC | KC:CONTAMINATION | The Product 1522SA evaluated for the contamination issue. The reported condition was observed on the physical sample. DHR Review cannot be performed as provided lot no. was not for 1522SA. DHR Review: Device History Record review cannot be performed as lot number 175631 does not belong to 1522SA product code. Trending Period: May 28, 2016 – June 19, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The sample evaluation was performed on returned samples. As per GS10000913 Grease (contamination) includes reaction plan to clean with clean cloth and isopropyl alcohol as part of regular production process. The probable root cause is human error and failure to follow normal cleaning procedure for the grease- contamination as per GS10000913 reaction plan to clean with clean cloth and isopropyl alcohol as part of regular production process. Also 1522SA is 2 QT. sharps disposal container which is not a sterile product. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk due to visual-cosmetics on the non-sterile product which is 2QT. Sharps Disposal container. Also as per complaint there is no patient involvement therefore does not pose any harm to the end user. | 0701870766 | 0 | Y | Japan | Closed | The customer reports there was brown line on the lid prior to use | Niti-on Co., Ltd | 0015877231 | 2017-05-29 | 2017-05-29 | 0 | 2017-05-28 | Prior to Use | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | Gatha Kothari | 42905 | 42905 | There are total 6 samples were submitted with this complaint. As from sample it was very minor grease contamination was observed on the parts. | Complete | #N/A
2017-05-29 | 0701871175-10 | 8836SA | UKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 3GA PHARMASTAR - TRANS WHITE | MITG - Sharps Disposal | 002927 | 8836SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8836SA evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record cannot be reviewed as lot number information is not provided. Trending Period: May 17, 2016 to June 09, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Also outside of Medtronic case may have been broken down to small quantity to ship to end user can cause missing absorbent pad issue. The absorbent pads are 2 packet of 5 each per packet (10 per case) to have one absorbent pad for each container as case quantity is 10 per case. There was no information in the complaint regarding quantity ordered by customer. The defective quantity listed as 4 eaches which means the customer did not receive full case of 10 eaches and case was broken down outside of Medtronic. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads should not be used. | 0701871175 | 0 | Y | United States | Closed | The customer states items should come with absorbent pad to help contain liquid content but the pads were not included in the package. | Medline Industries, Inc. | 0011846340 | 2017-05-17 | 2017-05-29 | 0 | 2017-05-17 | During use | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-05-31 | 0701875867-10 | 8985 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985 8 GAL CHEMO CONT HINGED YEL X10 | MITG - Chemo Protection Safety | 003209 | 8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8985 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: May 15, 2016 to June 09, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids product should not be used. | 0701875867 | 0 | Y | United States | Closed | The Customer is missing 1 lid for the containers. there was no damage to the case - case was sealed from manufacturer please have 1 lid sent for the container. | Medline Industries, Inc. | 0011846340 | 2017-05-15 | 2017-05-31 | 0 | 0 | Prior to Use | 2017-06-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-06 | 0701885051-10 | 31156916 | 22037959 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | FISHER 2 GAL CONTAINER; RED | MITG - Sharps Disposal | 002927 | 31156916 |  |  |  | RFR | CONTAINER ISSUE | AFC | KC:NEEDLE PUNCTURED CONTAINER | The Product 31156916 evaluated for the needle punctured container. The reported condition was observed in photo. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 167351 (10/01/2015 to 10/03/2015) and no internal reject found during production. Trending Period: June 05, 2016 to June 30, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family/lid type. Root Cause: The potential root cause is customer misuse by dropping sharps in perpendicular position. Also the container was autoclaved. So the sharps disposal content may have been moved during autoclaving and handling process resulted in the needle punctured through container. As per the photo in the attachment shows the needle is punctured through container perpendicularly. The DHR was reviewed and all inspection were well within specification. Also there is no information in the complaint regarding how full the container was. The sharps disposal containers are puncture resistance. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event due to customer misuse. The sharps were disposed perpendicular into the container instead of horizontal drop. There is no harm or injury reported for the complaint. | 0701885051 | 0 | Y | United States | Closed | The customer stated needle went through the side, post autoclaving. | Thermo Fisher Scientific | 0015890421 | 2017-06-05 | 2017-06-06 | 0 | 0 | Unknown | 2017-08-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | 0 | Janice Nevius | Gatha Kothari | 42916 | 42916 | No physical sample was available to review. The photo was attached in the complaint in attachment “FW Complaint#170190 Response required”. The photo shows needle punctured through container in the front side of container in the middle below the label. The lot number was found from the photo as 167351. The needle was punctured perpendicular to container. It is not very clear to determine if the container was above the fill line or not from photo evaluation. | Complete | #N/A
2017-06-09 | 0701894158-10 | 1181200777 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1181200777 NZ ONLY 12ML LUER LOCK SYRING | MITG - Needles & Syringes | 003208 | 1181200777 |  |  |  | RFR | PACKAGING | AFC | KC:NO FAILURE | Note Creator: Jeremy Roberts Note Date: 11/30/17 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was unable to be performed without a lot number. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Product analysis was unable to be performed. Without samples or pictures, the reported condition was not observed. An investigation was performed. The method of root cause analysis that was implemented was to perform an evaluation of the complaint file in combination with the device master record (DMR). The assessment evaluated potential causes for needles to be introduced based on product design. The investigation concluded product is not manufactured or supplied with bags. Sales unit of measure is 480 or 1000 pieces. Based on the information provided, the most probable cause originated within Fresenius Medical Care’s distribution channel where they are likely breaking down supplied product and placing it into bags of smaller units of measure. Based on the information available, trending analysis and investigation findings, additional actions are deemed unnecessary at this time. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0701894158 | 0 | Y | United States | Closed | The customer reported that when unloading a shipment, they noticed 10 individual needles inside a clear bag. | FRESENIUS MEDICAL CARE NA | 0012833563 | 2017-06-05 | 2017-06-09 | 0 | 0 | Unknown | 2017-12-02 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-12 | 0701896901-10 | 8938 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8938 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176259 (03/20/2017 to 03/22/2017) and no internal reject found during production. Trending Period: June 10, 2016 to June 22, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during packaging. The Medtronic case quantity for 8938 is 5 eaches/case. The complaint is for 1 each which shows the Medtronic case was broken down for further distribution to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701896901 | 0 | Y | United States | Closed | The customer states missing lids. | Medline Industries, Inc. | 0011846340 | 2017-06-10 | 2017-06-12 | 0 | 0 | Unknown | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-12 | 0701897993-10 | 8970Y |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8970Y evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: June 06, 2016 to June 22, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids packed prior to sealing a case. The Medtronic case quantity for 8970Y is 20 eaches/case. The complaint is for 20 eaches. There is no information available in complaint regarding total quantity ordered by the customer and if the product is received in Medtronic original packaging. As per customer complaint attachment mPXR-430865-2017-06-12 it stats that product return status as still in use. The product should not be used without lids. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701897993 | 0 | Y | Canada | Closed | The customer reports 20 missing lids from this order | Orillia Soldiers Memorial Hospital | 0001101349 | 2017-06-06 | 2017-06-14 | 0 | 2017-06-06 | During use | 2017-07-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Amin, Krishna PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Regulatory/Competent Authority ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: There were 20 lids missing from the order Below are the details for Product - Model No:8970Y, Description:8970Y YEL 2 GAL CONT ROTOR LID X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8970Y YEL 2 GAL CONT ROTOR LID X20 Model Number: 8970Y Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? Unknown Unknown, but will follow-up for more information SALESREP USER ID: roantm1 SALESREP EMP ID: 328283 SALESREP NAME: Roantree, Mike PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Roantree, Mike As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 328283 Roantree, Mike cannot find a BP in GCH ZINTCONT 328283 Roantree, Mike cannot find a BP in GCH ZFOLCONT 328283 Roantree, Mike cannot find a BP in GCH | Jill Hoffman | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-20 | 0701913934-10 | 8820- | 00175848 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 2GAL.PHARMASAFETY WAST CONT.20 | MITG - Sharps Disposal | 002927 | 8820- |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8820 evaluated for the missing absorbent pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record cannot be reviewed as lot number information is not provided. Trending Period: June 16, 2016 to June 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the container/lid style product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Also outside of Medtronic case may have been broken down to small quantity to ship to end user can cause missing absorbent pad issue. The absorbent pads are 4 packet of 5 each per packet (20 per case) to have one absorbent pad for each container as case quantity is 20 per case. There was no information in the complaint regarding quantity ordered by customer. The defective quantity listed as 1 each which means the customer did not receive full case of 20 eaches and case was broken down outside of Medtronic may resulted in missing absorbent pads. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads. There is no patient involvement, so there is no risk to patient safety or harm. | 0701913934 | 0 | Y | United States | Closed | The customer states Missing absorbent pads. | PATTERSON DENTAL SUPPLY | 0012856036 | 2017-06-16 | 2017-06-20 | 0 | 0 | Prior to Use | 2017-08-25 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-20 | 0701913938-10 | 8906 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 8906 YEL 1 QRT CONT PHLEBOTOMY X100 | MITG - Sharps Disposal | 002927 | 8906 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8906 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: June 14, 2016 to June 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaint reported for the product family which does not result in complaint trend yet. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids during packaging. The Medtronic case quantity for 8906 is 100 eaches/case. The complaint is for 60 eaches which shows the Medtronic case was broken down for further distribution in low unit of measure to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701913938 | 0 | Y | Canada | Closed | customer reports: : There are 60 missing lids from the case of sharps containers. | THE STEVENS COMPANY LIMITED | 0013267152 | 2017-06-15 | 2017-06-21 | 0 | 2017-06-14 | During use | 2017-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Bousquet, Jonathan PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Stevens - Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: 60 missing lids from case of sharps containers Below are the details for Product - Model No:8906, Description:8906 YEL 1 QRT CONT PHLEBOTOMY X100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8906 YEL 1 QRT CONT PHLEBOTOMY X100 Model Number: 8906 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No FACILITY NAME: Stevens FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 425 Railside Drive FACILITY CITY: Brampton FACILITY STATE CODE: ON FACILITY POSTAL CODE: L7A 0N8 FACILITY EMAIL: colleen.hewlin@stevens.ca SALESREP USER ID: molind1 SALESREP EMP ID: 252255 SALESREP NAME: Molino, Dayna PRIMARY CONTACT: Someone else PRIMARY CONTACT OTHER: Other PRIMARY CONTACT OTHER TEXT: Jonathan Bousquet jonathan.bousquet@medtronic.com COMPLAINT OWNER NAME: Molino, Dayna As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-21 | 0701915276-10 | V8536 | 00175968 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | VHA PLUS SSII 3GA CLR INRM 10 | MITG - Sharps Disposal | 002927 | V8536 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING CONTAINER | The Product V8536 evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175968 (04/18/2017 to 04/23/2017) and no internal reject found during production for the same issue. Trending Period: June 16, 2016 to June 29, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of containers prior to seal the case. The Medtronic case quantity for V8536 is 10 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701915276 | 0 | Y | United States | Closed | The customer states Received lids but no containers in sealed packaging. | Labette Health Physical Medicine | 0001308591 | 2017-06-16 | 2017-06-21 | 0 | 2017-06-16 | Prior to Use | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-21 | 0701915353-10 | 8938 | 00176763 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8938 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176763 (04/25/2017) and no internal reject found during production for the same issue. Trending Period: June 16, 2016 to June 29, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and nine other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The Medtronic case quantity for 8938 is 5 eaches/case. The complaint quantity is for 1 each. There is no other information in the complaint regarding if the product was in original Medtronic case or repackaged. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701915353 | 0 | Y | United States | Closed | The customer states no lids | HENRY SCHEIN INC DENVER | 0013101735 | 2017-06-16 | 2017-06-21 | 0 | 0 | Unknown | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-21 | 0701915353-20 | 8939 | 00177301 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8939 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 177301 (04/24/17 to 05/01/2017) and no internal reject found during production. Trending Period: June 16, 2016 to June 29, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and nine other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The Medtronic case quantity for 8939 is 5 eaches/case. There is no other information in the complaint regarding if the product was in original Medtronic case or repackaged. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701915353 | 0 | Y | United States | Closed | The customer states no lids | HENRY SCHEIN INC DENVER | 0013101735 | 2017-06-16 | 2017-06-21 | 0 | 0 | Unknown | 2017-07-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-22 | 0701918351-10 | 8881112599 | 17D04867X | COV LP LLC Crystal Lake DC | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | Existing/Previous Investigation | Existing/Previous Investigation | 0701918351 | 0 | Y | United States | Closed | customer reports: 6 cases of syringes had internal bags missing rubber bands. | AQUABILITI | 0012728353 | 2017-06-16 | 2017-06-23 | 0 | 2017-06-16 | During use | 2018-01-09 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Wayne Winton - Materials Manager ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Product still at Aquabiliti CONTACT NAME AND PHONE: Wayne Winton, Phone 615-833-2633 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event LOT# 17D04863X - 6 cases of syringes had internal bags missing rubber bands. Was the customer notified the product should be returned? No Please explain Return only with Crystal Lake QA approval SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires written investigation report. Also wants to return 6 cases - need Crystal Lake QA approval. | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-26 | 0701922307-10 | 8900SA | 00176458 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:LABEL ISSUE | The Product 8900SA evaluated for the label issue. The reported condition was observed on the physical sample which does not affect form, fit and functionality. DHR Review performed and met acceptance criteria. DHR Review: Device History Record reviewed for the lot number 176458 (04/26/17 to 05/12/17). There was no internal reject found during production. Trending Period: June 22, 2016 – July 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The 8900SA containers are molded, labeled and stacked during production process. There may be potentially label was not properly stick to the container and while stacking and packing one of the label corner slightly peeled off. The potential root cause is human error and failure to inspect label before packaging in to the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as all label information is well legible and except one corner label was slightly not adhered to the container. Also the issue was identified before use, there is no patient involved, therefore does not pose any harm to the end user. | 0701922307 | 0 | Y | Japan | Closed | customer reports: labeling came off of the package. | Niti-on Co., Ltd | 0015877231 | 2017-06-23 | 2017-06-26 | 0 | 2017-06-22 | Prior to Use | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Gatha Kothari | 42927 | 42927 | There are two samples received for evaluation. The both samples are from lot number 176458 for 8900SA. The bottom right corner slightly peeled off on both samples. In both samples product code, lot number, fill line and all other information are well legible. The reported condition observed but it will not affect form, fit and functionality of the part. The label was properly adhered to the container except slightly peeled on one corner. | Complete | #N/A
2017-06-26 | 0701923996-10 | 8605RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8605RC 5QT RCRA HAZARDOUS WASTE X14 | MITG - Sharps Disposal | 002927 | 8605RC |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:CRACKED CONTAINER | The Product 8605RC evaluated for the cracked container. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: June 26, 2016 to July 14, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The Medtronic case quantity is 14 eaches per case. As per the complaint reported for the 2 eaches which shows Medtronic case was broken into small quantity outside of Medtronic for distribution to end user. The potential root cause may be container have been cracked during shipping and handling due to external stress/forces. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product was inspected prior to use for any damage as per IFU. There is no harm or injury to end user as cracked container should not be used and must be replaced. | 0701923996 | 0 | Y | United States | Closed | The customer states containers were cracked. | Redaway LLC | 0016281506 | 2017-06-26 | 2017-06-26 | 0 | 2017-06-26 | Prior to Use | 2017-09-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Asked but Unknown | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-06-27 | 0701927079-10 | 8970Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:CRACKED CONTAINER | The Product 8970Y evaluated for the cracked container. The reported condition was observed in the photos available for product analysis. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: June 13, 2016 to July 14, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed one internal inspection reject and one other complaint reported for the product code. Root Cause: The lot number is unknown so DHR cannot be reviewed. As lot number was unknown it cannot be determined when the product was manufactured. The packaging improvement was implemented as per CDP-0005979 in June, 2016 per CAPA2015-01007800 for 8970Y. As there is no lot number available it cannot be determined whether the product is made prior to the June 2016. The potential root cause may be container have been cracked during shipping and handling due to external stress/forces manufactured prior to June 2016 packaging improvement change. Also it was noted in the complaint there was no safety or harm to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. | 0701927079 | 0 | Y | Canada | Closed | The customer reports a big hole/crack along the bottom of the container. They also received 3 new ones that were broken when I removed them from the box they were delivered in. | Guildford Public Health Centre | 0016136021 | 2017-06-27 | 2017-06-28 | 0 | 2017-06-13 | During use | 2017-08-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: The container was picked up by me from the counter top because it was full, so I wanted to dispose of it. I heard something glass-like fall on the floor and I looked down and saw a used syringe. I lifted the sharps container and saw a big hole/crack along the bottom. I have also received 3 new ones that were broken when I removed them from the box they were delivered in. Below are the details for Product - Model No:8970Y, Description:8970Y YEL 2 GAL CONT ROTOR LID X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8970Y YEL 2 GAL CONT ROTOR LID X20 Model Number: 8970Y Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? During use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Vandeleur PHYSICIAN FIRST NAME: Rosalie PHYSICIAN CLINIC: Guildford Public Health Centre PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Surrey PHYSICIAN STATE CODE: BC PHYSICIAN DAY PHONE: 604-587-4760 PHYSICIAN EMAIL: rosalie.vandeleur@fraserhealth.ca For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes FACILITY NAME: Guildford Public Health FACILITY COUNTRY CODE: CAN FACILITY CITY: Surrey FACILITY STATE CODE: BC SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Dahlbeck, Michelle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZFOLCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH | Daniel Tamul | Gatha Kothari | 42930 | 42930 | There photos were attached in the complaint. The reported condition was observed in the photos. (101283-Sharps1.JPG and 101284-Sharps2.JPG) The container was cracked in the bottom.. | Complete | Yes
2017-06-28 | 0701928868-10 | 8985 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985 8 GAL CHEMO CONT HINGED YEL X10 | MITG - Chemo Protection Safety | 003209 | 8985 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING COMPONENTS | The Product 8985 evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: June 23, 2016 to July 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of containers/lids prior to seal the case. The Medtronic case quantity for 8985 is 10 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701928868 | 0 | Y | United States | Closed | The customer states missing 1 each. | Seneca Medical Inc | 0016251749 | 2017-06-28 | 2017-06-28 | 0 | 2017-06-23 | Unknown | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-06-29 | 0701932743-10 | UNK NS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | NEEDLES + SYRINGES UNKNOWN | MITG - Needles & Syringes | 003208 | UNK NS |  |  |  | RFR | BURN(S) | AFC | KC:UNIDENTIFIED LIQUID | Note Creator: Jeremy Roberts Note Date: 8/17/17 Note Description: Addition of Investigation Summary. Investigation Summary: Incident Description: Customer reports: Condensation inside syringe casing... when opened, the liquid has burned the skin of 3 individuals. Two subsequent complaints have been entered; one for each individual case. PE’s: 0701932743, 0701932744 and 0701709213. All complaints shall be investigated as one complaint. Product Analysis: As of 8/15/17 samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and investigation may be performed. Based on the reported information and pictures provided the following was noted: 1. GCH attachment “Condensate”, shows a hard pack luer lock syringe that has been removed from its primary packaging and reassembled/closed. The reported condition of condensation is observed in the sleeve. The Cap is shown on top of the Sleeve with a Luer Lock Syringe in the users hand. a. The only hard pack luer lock syringe made in Crystal Lake is 8881512878. b. 8881512878 is a sterile, hard pack, single use product. 2. GCH attachments “Thumb” and “thumb 2” show two separate thumbs with the incident description indicating the liquid has burned the skin of 3 individuals. a. This is an indication that the sterile single use syringe was used without gloves. It is not clear from the information provided if the same syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size for part 8881512878 is 360 cases and production line 23 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a process walk while performing a Six M assessment to identify areas where liquid is introduced into the manufacturing process that could make its way inside the molded sleeve. The assessment evaluated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Significant - Controlled Manufacturing Cleanliness GS20002641 does not have a quiz associated with the training item. Inclusion of a quiz will increase training effectiveness. GS20002641 outlines pre-cleaning activities requiring parts be removed from the manufacturing equipment. In the event that product remains in the Aidlin during periodic cleaning activities, cleaning agents can make their way into the sleeve and appear as condensation.) Material Causes: -Cleaning agent (Not Significant - Vesta Syde SQ Disinfectant is used in the manufacturing environment for cleaning on a periodic basis. Processes are in place to halt Manufacturing and protect product during cleaning activities. Burns or irritation to skin may occur if not properly used.) - Silicone (Not Significant – DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512878 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. In the event of a malfunction, silicone could be introduced into the sleeve and appear as condensation. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Burns or irritation to skin is unlikely). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and is required to be used during manufacturing). Method Insufficient Cleaning Process (Significant – Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. In the event solution is not properly diluted skin burns or irritation may occur.) Machine Causes: N/A: The only location within the manufacturing process where liquid is introduced with the possibility of entering the sleeve is during assembly where silicone is added to the syringe barrel. Reference material section above. Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: In the event of a failure during the manufacturing process, the most probable root cause is insufficient mixing of Vesta Syde SQ Disinfectant and Operator Training to GS20002641 Controlled Manufacturing Cleanliness GS20002641. It is not clear from the information provided if the syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Unidentified Liquid) and product family was performed. No other incidences of this nature for this product family were received in the trending period from 8.15.17 – 03.02.16. Zero internal rejects of this nature for this product family were received in the trending period from 8.15.17 – 12.2.16. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. In accordance with Medical Safety Review GCH transaction ID 226601208, Severity for this incident is a level 1. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. ComplianceWire training item CRY-GS20002641 Version 206.0.1 went effective 8/15/17 to active the quiz and initiate retraining for all affected employees. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure particulate and extraneous matter. | 0701932743 | 0 | Y | United States | Closed | customer reports: Condensation was inside the syringe casing and when opened, the liquid has burned the skin of 3 individuals. | SUMMIT VETERINARY REFERRAL CTR | 0013150347 | 2017-03-02 | 2017-06-30 | 0 | 2017-03-01 | During use | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Non-Serious Illness Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | Yes
2017-06-30 | 0701932744-10 | UNK NS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | NEEDLES + SYRINGES UNKNOWN | MITG - Needles & Syringes | 003208 | UNK NS |  |  |  | RFR | BURN(S) | AFC | KC:UNIDENTIFIED LIQUID | Note Creator: Jeremy Roberts Note Date: 8/17/17 Note Description: Addition of Investigation Summary. Investigation Summary: Incident Description: Customer reports: Condensation inside syringe casing... when opened, the liquid has burned the skin of 3 individuals. Two subsequent complaints have been entered; one for each individual case. PE’s: 0701932743, 0701932744 and 0701709213. All complaints shall be investigated as one complaint. Product Analysis: As of 8/15/17 samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and investigation may be performed. Based on the reported information and pictures provided the following was noted: 1. GCH attachment “Condensate”, shows a hard pack luer lock syringe that has been removed from its primary packaging and reassembled/closed. The reported condition of condensation is observed in the sleeve. The Cap is shown on top of the Sleeve with a Luer Lock Syringe in the users hand. a. The only hard pack luer lock syringe made in Crystal Lake is 8881512878. b. 8881512878 is a sterile, hard pack, single use product. 2. GCH attachments “Thumb” and “thumb 2” show two separate thumbs with the incident description indicating the liquid has burned the skin of 3 individuals. a. This is an indication that the sterile single use syringe was used without gloves. It is not clear from the information provided if the same syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size for part 8881512878 is 360 cases and production line 23 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a process walk while performing a Six M assessment to identify areas where liquid is introduced into the manufacturing process that could make its way inside the molded sleeve. The assessment evaluated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Significant - Controlled Manufacturing Cleanliness GS20002641 does not have a quiz associated with the training item. Inclusion of a quiz will increase training effectiveness. GS20002641 outlines pre-cleaning activities requiring parts be removed from the manufacturing equipment. In the event that product remains in the Aidlin during periodic cleaning activities, cleaning agents can make their way into the sleeve and appear as condensation.) Material Causes: -Cleaning agent (Not Significant - Vesta Syde SQ Disinfectant is used in the manufacturing environment for cleaning on a periodic basis. Processes are in place to halt Manufacturing and protect product during cleaning activities. Burns or irritation to skin may occur if not properly used.) - Silicone (Not Significant – DOW CORNING Silicone Number 360, 12500 centi stokes is a component on the bill of material for product code 8881512878 and is expected to be inside the syringe. During the manufacturing process, a small amount of silicone is sprayed into the syringe barrel prior to the addition of the plunger. In the event of a malfunction, silicone could be introduced into the sleeve and appear as condensation. Silicone is added to the syringe to facilitate actuation of the plunger. This material possesses characteristics which include low chemical reactivity, low toxicity, and it does not support microbiological growth. Burns or irritation to skin is unlikely). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and is required to be used during manufacturing). Method Insufficient Cleaning Process (Significant – Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. In the event solution is not properly diluted skin burns or irritation may occur.) Machine Causes: N/A: The only location within the manufacturing process where liquid is introduced with the possibility of entering the sleeve is during assembly where silicone is added to the syringe barrel. Reference material section above. Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: In the event of a failure during the manufacturing process, the most probable root cause is insufficient mixing of Vesta Syde SQ Disinfectant and Operator Training to GS20002641 Controlled Manufacturing Cleanliness GS20002641. It is not clear from the information provided if the syringe was re-used. As indicated on the instructions for use provided, it is advised that syringes not be used if opened or damaged. A tamper evident seal is provided to determine if the syringe has been opened. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Unidentified Liquid) and product family was performed. No other incidences of this nature for this product family were received in the trending period from 8.15.17 – 03.02.16. Zero internal rejects of this nature for this product family were received in the trending period from 8.15.17 – 12.2.16. Formal Investigation: Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. In accordance with Medical Safety Review GCH transaction ID 226601208, Severity for this incident is a level 1. Trending triggers are established to determine when use of the Quality Signal Evaluation and Action System Decision Tool should be used. At this point, a trend has not been established. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Corrective and Preventive Action: Change order QC021634 was completed 3/24/17 to identify which department is responsible for Vesta Syde SQ solution preparation and provide instructions on how to properly mix/dilute the cleaning agent. ComplianceWire training item CRY-GS20002641 Version 206.0.1 went effective 8/15/17 to active the quiz and initiate retraining for all affected employees. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure particulate and extraneous matter. | 0701932744 | 0 | Y | United States | Closed | customer reports: Condensation was inside the syringe casing and when opened, the liquid has burned the skin of 3 individuals. | SUMMIT VETERINARY REFERRAL CTR | 0013150347 | 2017-03-02 | 2017-06-30 | 0 | 2017-03-01 | During use | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Non-Serious Illness Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | Yes
2017-06-30 | 0701934263-10 | 8790 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 2 GA NON-INFECTIOUS GLASS CONT | MITG - Sharps Disposal | 002927 | 8790 |  |  |  | RFR | LABELING ISSUE | AFC | KC:MISSING OR DAMAGED LABEL | The Product 8790 evaluated for the label issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: Feb 23, 2016 – July 13, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: For root cause analysis the complaint attachment “{5718321} 601934263” was reviewed. As per attachment it states “The customer ordered 8 cases of the item, however on 1 of the case there was no label.” So the event description should be specific stating missing overshipper label. Based on this information the probable root cause is human error and failure to verify overshipper label presence on the case before placing case on the skid. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as overshipper label is additional label for product information. There is no evidence of any label integrity affected for product labeling from complaint information. | 0701934263 | 0 | Y | United States | Closed | The customer states Did not have label on the product. | VWR INTERNATIONAL | 0013143732 | 2017-02-24 | 2017-06-30 | 0 | 2017-02-23 | Unknown | 2017-07-27 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-04 | 0701939124-10 | 1522SA | 00173931 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | CONTAMINATION | AFC | KC:CONTAMINATION | The Product 1522SA evaluated for the contamination issue. The reported condition was observed on the photos provided. DHR Review was performed and met the specification. DHR Review: Device History Record reviewed for the lot number 173931 (12/05/16 to 12/12/16). There was no internal reject found during production. Trending Period: Feb 14, 2016 – July 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection rejects and no other complaint for the product family. Root Cause: This issue has been determined to be an isolated event with minimal risk as contaminated product should not be used. Before use container/lid needs to be inspected for any abnormality. The potential root cause may be failure to inspect product for contamination prior to seal the finished good case. Also one of the photo shown the non-Medtronic Overshipper case which may be directly related to repackaging of the product outside of Medtronic which may resulted in to the different packaging of product and also contamination of product during handling due to human error. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as contaminated product should not be used. The product should be inspected prior to use for any abnormality. | 0701939124 | 0 | Y | United States | Closed | The customer states that two of the six boxes were packaged differently than past. inside one of the oddly shaped boxes, most of the sharps had sticky residue all over them. | Thermo Fisher Scientific | 0015890421 | 2017-02-14 | 2017-07-04 | 0 | 0 | Unknown | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Gatha Kothari | 42942 | 42942 | The sample photos were submitted with this complaint. The photo 20170210_105633 is shown Medtronic Packaging box per specification. The photo 20170210_105652 has photo of Overshipper box looks shorter than Medtronic box. The other photos shown have some contamination on the product. It cannot be determined what kind of contamination as only photos are available for review. | Complete | #N/A
2017-07-04 | 0701939184-10 | 8989 | 00174559 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174559 (01/18/17 - 01/22/17) and no internal reject found during production for the same issue. Trending Period: July 04, 2016 to July 25, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and ten other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701939184 | 0 | Y | United States | Closed | The Customer states missing the lids. | ATLANTA CANCER CARE PERIMETER | 0012855903 | 2017-02-16 | 2017-07-04 | 0 | 2017-02-16 | Unknown | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-05 | 0701940001-10 | 8881112059 | 15G09363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | LUER ISSUE | AFC | KC:DAMAGED COMPONENTS | Note Creator: Jeremy Roberts Note Date: 9/12/17 Note Description: Addition of Investigation. Incident Description: Defect in the Syringe Connector, Visual observation found the syringe to be bent or crimped on the slip fit tip, with two impressions opposite each other on the shaft of the tip. There was some transverse scratching on the syringe barrel, and the printing on the syringe was deformed and incomplete. There were missing or inaccurate graduation marks. Product Analysis: 1 Sample was sent for evaluation. It was not in its original supplied packaging. The reported conditions were observed. Based on the sample review it is clear in that a jam occurred during the printing operation which damaged the barrel and cause skewed printing. DHR Review: A review of the device history record was completed for lot 15G09363X. 180 Cases (90,000 pieces) were produced and released using production line 7 on 8/1/2015. The DHR was review showed no abnormal process conditions present during the manufacturing of this product that would lead to the reported conditions. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process, inclusive of running using validated process settings. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The assessment eval uated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Not Significant – Operators are trained to perform required functions in accordance with instructions provided in the design specification). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and is used during manufacturing to support lot release). Material Causes: - Ink and Resin: (Not Significant – Materials are a supplied component confirmed to meet acceptance criteria upon receipt as required). Method Printing Process (Significant - In the event of incorrect start-up after a machine jam (damaged component), non-conforming parts may be passed to downstream operations.) Machine Causes: Printer (Not Significant – PS10002933 Process Control SOP is in place and being used to support production. 1 observed defect is within the validated range of the equipment, meeting acceptance criteria). Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: The most probable root cause is that a printer machine jam damaging the component and the process was incorrectly started back up following the jam. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Damaged Product/Components) and product family was performed. 1 other incidents of this nature for this product family was received in the trending period from 7.5.16 – 9.12.17. Zero internal NCR rejects of this nature for this product family were received in the trending period from 5.5.17 – 9.12.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time. However, it is noted that DC00045513 was released on 4/15/16 (after production of the reported lot) whereby Standard Work was put in place for Line 7 Printer Start Up operations. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Review: RE00035287 Rev. 0 pFMEA for Syringe Pad Printer was reviewed. It was confirmed that damaged component/barrel is included with appropriate controls in place. | 0701940001 | 0 | Y | United States | Closed | The customer reports the defect is in the syringe connector. Visual observation found the syringe to be bent or crimped on the slip fit tip, with two impressions opposite each other on the shaft of the tip. There was some transverse scratching on the syringe barrel, and the printing on the syringe was deformed and incomplete. There were missing or inaccurate graduation marks. | BARD SHANNON LIMITED | 0012818702 | 2017-07-05 | 2017-07-06 | 0 | 2017-07-05 | During use | 2017-10-05 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Serial/Lot No:15G09363X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 15G09363X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event Defect in the Syringe Connector, Visual observation found the syringe to be bent or crimped on the slip fit tip, with two impressions opposite each other on the shaft of the tip. There was some transverse scratching on the syringe barrel, and the printing on the syringe was deformed and incomplete. Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: Defect in the Syringe Connector, Visual observation found the syringe to be bent or crimped on the slip fit tip, with two impressions opposite each other on the shaft of the tip. There was some transverse scratching on the syringe barrel, and the printing on the syringe was deformed and incomplete. | Janice Nevius | Jeremy Roberts | 42990 | 42990 | Note Creator: Jeremy Roberts Note Date: 9/12/17 Note Description: Addition of Analysis. Product Analysis: 1 Sample was sent for evaluation. It was not in its original supplied packaging. The reported conditions were observed. Based on the sample review it is clear in that a jam occurred during the printing operation which damaged the barrel and cause skewed printing. | Complete | Yes
2017-07-05 | 0701940281-10 | 85321 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85321 CLEAR 8 QUART CONT OPEN LIDX10 | MITG - Sharps Disposal | 002927 | 85321 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 85321 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 05, 2016 to July 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the product family/different style lids. There is no other complaint for 85321 product code. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701940281 | 0 | Y | United States | Closed | The customer reports 2 lids for the sharps containers are missing from the 10 each case that was ordered. | N CENTRAL HEALTH CARE FACILITY | 0013241200 | 2017-07-05 | 2017-07-06 | 0 | 2017-07-05 | Prior to Use | 2017-08-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Bolt, Alison PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: 2 lids for the sharps containers are missing from the 10 each case that was ordered. Below are the details for Product - Model No:85321, Description:In Room Sharps Container, Always Open Lid, Clear IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for In Room Sharps Container, Always Open Lid, Clear Model Number: 85321 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-06 | 0701941342-10 | 89651 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 89651 2GAL RED CONT H DROP CLR LIDX20 | MITG - Sharps Disposal | 002927 | 89651 |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | KC:MISSING WELD | The Product 89651 evaluated for the missing weld. The reported condition was observed during sample evaluation. The cowl and lid sample was missing weld operation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 05, 2016 to Aug 01, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to follow lid assembly welding operation step during manufacturing before packaging lids in to the case. The lot number is unknown so DHR review cannot be performed. The Standard work for assembly and inspection procedure were reviewed and found adequate for manufacturing process. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event as unassembled lid should not be used. There is no known impact or consequence to patient or end user safety. | 0701941342 | 0 | Y | Japan | Closed | The customer reports product lid was malformed and it came off from the body easily. No patient harm. | Niti-on Co., Ltd | 0015877231 | 2017-07-06 | 2017-07-06 | 0 | 2017-07-05 | Prior to Use | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | Gatha Kothari | 42944 | 42944 | The sample received is only lid WP8965N lid. The lid and cowl component were received without welding operation. As from sample evaluation it was observed that lid assembly steps was missed during manufacturing operation. The cowl and lid missed the weld operation for the assembly. | Complete | Yes
2017-07-11 | 0701950417-10 | 8932PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8932PG2 12 GAL RED PG2 HINGE LID X10 | MITG - Sharps Disposal | 002927 | 8932PG2 |  |  |  | RFR | LID ISSUE | AFC | KC:MISSING LIDS | The Product 8932PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 09, 2016 to Aug 04, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids. The case quantity for 8932PG2 is 10 eaches/case. The complaint is for 5 each which shows the case was broken down for further distribution to end user. There is no information available in complaint regarding if customer received product in original Medtronic case or repack from third party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701950417 | 0 | Y | United States | Closed | the customer reports missing 5 lids. | MEDLINE | 0012894606 | 2017-07-09 | 2017-07-11 | 0 | 2017-07-09 | Unknown | 2017-08-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Stacey Olsen | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-11 | 0701950888-10 | 8605RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8605RC 5QT RCRA HAZARDOUS WASTE X14 | MITG - Sharps Disposal | 002927 | 8605RC |  |  |  | RFR | CRACKED CONTAINER | AFC | KC:CRACKED CONTAINER | The Product 8605RC evaluated for the cracked container. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number is unknown. DHR Review: A Device History Record was not reviewed as lot number was unknown. Trending Period: July 11, 2016 to Aug 04, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. The same complaints reported twice with different Product Event numbers from the same customer. So it will be consider as one complaint and will not count towards trending. (PE: 0701950888 and PE: 701923996) Root Cause: The Medtronic case quantity is 14 eaches per case. As per the complaint reported for the 2 eaches which shows Medtronic case was broken into small quantity outside of Medtronic for distribution to end user. The potential root cause may be container have been cracked during shipping and handling due to external stress/forces. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product was inspected prior to use for any damage as per IFU. There is no harm or injury to end user as cracked container should not be used and must be replaced. | 0701950888 | 0 | Y | United States | Closed | The customer states that the containers were cracked.-###-From PE-701923996-The customer states containers were cracked. | Redaway LLC | 0016281506 | 2017-07-11 | 2017-07-11 | 0 | 0 | Unknown | 2017-09-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | Yes
2017-07-20 | 0701969836-10 | 85120Y | 201705310012 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85120Y YEL 5 QRT CONT ALWAYS OPEN LID | MITG - Sharps Disposal | 002927 | 85120Y |  |  |  | RFR | PACKAGING | AFC | KC:NO FAILURE | The Product 85120Y evaluated for “no failure”. The reported condition cannot be observed as no photo or physical sample available for evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. The lot number provided is not a valid lot number per Cardinal Health. Trending Period: July 17, 2016 to Aug 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The root cause is user error to follow IFU (Instruction for Use) for final lock. 85120Y does not have clip by design. The 85120Y product has “Always open lid”. The lid design does not have clip for final lock. For final lock –door needs to be pushed down to secure by past locks in lid. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: Based on the above information, there was no failure found with the product per complaint. The product is manufactured as per intended use. | 0701969836 | 0 | Y | Canada | Closed | Customer reports that locking clips are missing. | CARDINAL HEALTH CANADA INC RDP | 0012702795 | 2017-07-18 | 2017-07-21 | 0 | 2017-07-17 | Unknown | 2017-09-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-20 | 0701970536-10 | UNK SCS | 00173677 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | SHARPS CONTAINERS UNKNOWN | MITG - Sharps Disposal | 002927 | UNK SCS |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 80-8705 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 173677 (03/09/17 to 03/13/2017) and no internal reject found during production. From the lot number listed in complaint; Product ID was obtained. Trending Period: July 20, 2016 to Sept 06, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the required amount of lids prior to packaging/handling. The case quantity for 80-8705 is 10 eaches/case. From the complaint description it states “25 missing lids” but there is no information about how many cases were received by customer. As from the information provided the probable root cause is handling of product from distribution to the customer. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701970536 | 0 | Y | United States | Closed | According to the reporter, on July 20, 2017, during pre operation, there were only 4 pieces from stock and found only 5 lids per case. The customer needed 25 lids for the item. There was no patient involvement. | 0 | 0 | 2017-07-20 | 2017-08-31 | 0 | 2017-07-20 | Prior to Use | 2017-09-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | No Product found for LotSerial# 00173677 FIRST NOTIFIED NAME: Platenyk, Kyle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: McKesson - Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: McKesson pulled 4cs of lot 00173677 from stock. There are only 5 lids per case. McKesson needs 25 lids for item 80-8705 . Below are the details for Product - Serial/Lot No:00173677 , Description:80-8705 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 80-8705 Model Number: Not specified Serial or Lot Number: 00173677 ------------------------------------------------------------------------ -------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No SALESREP USER ID: platek1 SALESREP EMP ID: 326455 SALESREP NAME: Platenyk, Kyle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Platenyk, Kyle | Rose Matthews | Gatha Kothari | 42984 | 42984 | No samples were submitted with this complaint. The reported condition stated by the customer could not be determined as samples or photos were not available for evaluation. | Complete | #N/A
2017-07-24 | 0701975416-10 | 8881112059 | 16L05663X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:BROKEN LUER TIP | Note Creator: Jeremy Roberts Note Date: 10/12/17 Note Description: Addition of Investigation Summary. Incident Description: 1 Syringe missing luer slip tip. Attachment “104267-Covidien FYI SCAR003243.docx” goes on to further state “While performing EOL on Job# 5216667 of 24pcs, QA found 1/20 with missing male slip on PN 10127503, LN 630140 / SLN 16L05663X.” Product Analysis: One syringe was returned for evaluation. The syringe was not in originally supplied packaging. The sample was confirmed to be from Mold 58115 cavity 41B. Upon visual inspection of the syringe subassembly, the barrel luer tip was observed to be broken off with a small portion of the luer base remaining. Magnified examination also observed a hairline crack of approximately 0.007” long on the barrel shoulder. DHR Review: A review of the device history record was completed for lot 16L05663X. The lot released 288 cases (144,000 pieces) on 4/14/17 using production line 7. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes. Significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: - Training - Improper setting of sensors or air (Not Significant – improper setting would be easily detected through line/machine outputs) - Clearing of Aidlin jams (Not Significant – Standard work is in place and posted at point of use for managing Jams). - Excess or angular stress on tip (Potentially Significant – If the user puts excessive or even angular stress on the tip during further processing/handling the tip may break). Material Causes: -Cold parts (Not Significant – during startups, cold parts are introduced with no observed impact) -Resin (Not significant – resin has not changed, all lots are certified for compliance prior to use) - Mating Medical Device Defective (Potentially significant – if the customer is applying a defective mating device this could damage the tip and cause premature failure/breakage). Method Causes: - Preventative/Maintenance – Conveyors (Not Significant – Preventative maintenance is in place to insure conveyors are operating properly). Measurement Causes: - Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). - Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). - Validated process parameters (Not significant – Parameters are confirmed as part of lot acceptance activities) Machine Causes: - Speeds and feeds (Not significant – speeds and feeds are kept within process control ranges where applicable. Other changes would be detected through line outputs) - Aidlin material feed (Potentially Significant – RAM validation along with continuous process control confirms common cause material feed variation is within expected operating range. This variation has the potential to create breakage). - Printer alignment (Not significant - Improper alignment was not observed) -Hopper/Conveyer belt clearance (Not Significant – The conveyor belts were reviewed and confirmed to have adequate clearance) - Transfer tube alignment/barrel drop station on RAM (Not Significant – Barrel drop station is set in accordance with established standard work). Environment Causes: -Humidity and Temperature – (Not significant – humidity and temperature are controlled in the CME). Root Cause: Root cause investigation could not confirm a manufacturing defect as this may occur if excess or angular stress is applied during further processing (i.e. If the user puts excessive or even angular stress on the tip during further processing/handling the tip may break ¦if the customer is applying a defective mating device this could damage the tip and cause premature failure/breakage). AngioDynamics further processes the Syringe. In the event a defect was created during the manufacturing process, the most probable root cause is common cause variation during Aidlin equipment handling. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Broken Luer Tip) and product family was performed. Zero other incidents of this nature for this product family was received in the trending period from 7.20.16 – 10.12.17. Zero internal rejects of this nature for this product family were received in the trending period from 4.20.17 – 10.12.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, a formal investigation is deemed unnecessary at this time. Based on the risk and potential of occurring during manufacturing, the following preventative actions have been taken: (1) Work order CL-2653039 was approved to shut down assembly equipment when low level sensors are activated for a specified time frame. Low level sensor activation occurs when Aidlin material feed variation is experienced. (2) Change order EC091212 institutes a line challenge checklist to ensure relevant systems are setup and operating as expected. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure for damage non-functional and associated controls. Line challenge check list is in support of existing pFMEA. | 0701975416 | 0 | Y | United States | Closed | The customer reports one syringe missing luer slip tip. | ANGIODYNAMICS | 0012721304 | 2017-07-21 | 2017-07-25 | 0 | 2017-07-20 | During use | 2017-10-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: QA Coordinator Danielle Abbenante ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: 1 Syringe missing luer slip tip. Below are the details for Product - Model No:8881112059, Serial/Lot No:16L05663X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 16L05663X ------------------------------------------------------------------------ ----------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? No Please explain Just 1 syringe and sent in as QA sample. SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires close out letter. Sample to be forwarded to Decontam Lab. Contact Danielle Abbenante, Phone # 518-792-4112 x 4254. E-Mail: dabbenante@angiodynamics.com | Shanika Hodge | Jeremy Roberts | 43020 | 43020 | Note Creator: Jeremy Roberts Note Date: 10/12/17 Note Description: Addition of Product Analysis. Product Analysis: One syringe was returned for evaluation. The syringe was not in originally supplied packaging. The sample was confirmed to be from Mold 58115 cavity 41B. Upon visual inspection of the syringe subassembly, the barrel luer tip was observed to be broken off with a small portion of the luer base remaining. Magnified examination also observed a hairline crack of approximately 0.007” long on the barrel shoulder. | Complete | #N/A
2017-07-24 | 0701975437-10 | 1522SA | 17E06563 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 1522SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E06563 (05/31/17 to 06/14/17) and no internal reject found during production for the same issue. Trending Period: July 24, 2016 to Aug 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family/different style lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 1522SA is 60 eaches/case. There is no information in the complaint regarding if the customer received product in original Cardinal Health overshipper case with 60 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701975437 | 0 | Y | United States | Closed | The customer reports 4 lids were missing from the case of product. | DIK FULFILLMENT SERVICES | 0013114437 | 2017-07-21 | 2017-07-25 | 0 | 2017-07-21 | During use | 2017-09-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:1522SA, Serial/Lot No:17E06563, Description:1522SA 2 QRT RED CONTAINER PHLEBOTOMY IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 1522SA 2 QRT RED CONTAINER PHLEBOTOMY Model Number: 1522SA Serial or Lot Number: 17E06563 ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth ADDITIONAL COMMENTS ABOUT THE EVENT: 4 lids were missing from the case of product. Please re-ship 4 lids | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-25 | 0701977428-10 | 8881112985 | 16B00563X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 10ML STERILE WATER SYR BULK NS | MITG - Needles & Syringes | 003208 | 8881112985 |  |  |  | RFR | LUER ISSUE | AFC | KC:DEFORMED/WARPED | Note Creator: Jeremy Roberts Note Date: 8/22/17 Note Description: Addition of Investigation Summary. Incident Description: As reported event description: The tip of the syringe was damaged. Note: 8881112985 is provided to Nurse Assist as a bulk non-sterile syringe for further processing. Nurse Assist fills the supplied syringe with sterile water, places a tip cap on the syringe and then packages it for sale to market. Product Analysis: Two syringes were returned for evaluation (sample 1 and sample 2) along with a non-medtronic catheter. Neither syringe is in originally supplied packaging. 1: Sample 1 was confirmed to be from Mold 58020 cavity 70. The sample was visually inspected for defects. The luer tip is intact with no abnormalities observed and was confirmed to meet ISO 594 luer taper requirements (Gage ID: SR-734, Calibration due 5/16/18). The syringe barrel shoulder was observed to be concaved forcing the rubber tip and plunger rod away from the fully seated position. This deformation is commonly seen during barrel molding and is due to material build-up on mold core pins which causes the syringe barrel to stick during mold ejection. 2: Sample 2 was confirmed to be from Mold 58020 cavity 68. Upon visual inspection of the syringe subassembly, the barrel luer tip was observed to be broken off with a small portion of the luer base remaining. The broken luer tip was observed to be stuck inside the catheter. Magnified examination confirmed that an angular force was applied to the luer which caused it to snap off and subsequently create a crack on the barrel shoulder extending to the barrel side wall. The broken luer was removed from the catheter and inspected to confirm conformance to the ISO 594 standard (Gage ID: SR-734, Calibration due 5/16/18). DHR Review: A review of the device history records were completed for lots 16B00563X and 16D02163X. Each lot produced and released 700 cases (700,000 pieces) on 2/17/16 and 4/17/16 respectively using production line 7. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process, inclusive of running using validated process settings. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes specifically for the broken luer tip. Cause of luer tip damage/deformation was identified under product analysis. Significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: - Training - Improper setting of sensors or air (Not Significant – improper setting would be easily detected through line/machine outputs) - Clearing of jams (Not Significant – jammed parts are discarded in accordance with established standard work) - Excess or angular stress on tip (Potentially Significant – If the user puts excessive or even angular stress on the tip during use the tip may break) Material Causes: -Cold parts (Not Significant – during startups, cold parts are introduced with no observed impact) -Resin (Not significant – resin has not changed, all lots are certified for compliance prior to use) - Mating Medical Device Defective (i.e. Catheter or Tip cap connector defect or malfunction - Potentially significant – if the mating device is defective it could damage the tip and cause premature failure/breakage) Method Causes: - Preventative maintenance – Conveyors (Not Significant – Preventative maintenance is in place to insure conveyors are operating properly). - Speeds and feeds (Not significant – most speeds and feeds are kept within validated ranges. Other changes would be detected through line outputs) - Tip Cap Application Process (Potentially Significant – If the secondary manufacturer puts excessive or angular force on the luer tip during tip cap application the tip may weaken or break. How much force is used when applying the tip cap?). Measurement Causes: - Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). - Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). - Validated process parameters (Not significant – Parameters are confirmed as part of lot acceptance activities) Machine Causes: - Printer alignment (Not significant - Improper alignment was not observed) -Hopper/Conveyer belt clearance (Not Significant – The conveyor belts were reviewed and confirmed to have adequate clearance) - Transfer tube alignment/barrel drop station on RAM (Not Significant – Barrel drop station is set in accordance with established standard work) Environment Causes: -Humidity and Temperature – (Not significant – humidity and temperature are controlled in the CME). Root Cause: Root cause investigation could not confirm a manufacturing defect as this may occur if excess or angular stress is placed on the tip during mating or a malfunctioning mating device damaged the tip causing premature failure/breakage. Trend Analysis – Sample 1 (Deformation): GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Deformation) and product family was performed. one other incident of this nature for this product family was received in the trending period from 7.25.16 – 8.22.17. Zero internal rejects of this nature for this product family were received in the trending period from 5.25.17 – 8.22.17. Formal Investigation, Corrective and Preventive Action – Sample 1 (Deformation): Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the confirmed deformation condition for this product family presents a low risk level. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Trend Analysis – Sample 2 (Component Broke in Use): GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Component Broke in Use) and product family was performed. Zero other incidents of this nature for this product family was received in the trending period from 7.25.16 – 8.22.17. Zero internal rejects of this nature for this product family were received in the trending period from 5.25.17 – 8.22.17. Formal Investigation, Corrective and Preventive Action – Sample 2 (Component Broke in Use): Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. The reported condition could not confirm a manufacturing defect for this complaint. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: Sample 1: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed inclusion of failure for deformation and associated controls. Sample 2: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure for damage non-functional and associated controls. | 0701977428 | 0 | Y | United States | Closed | The customer reports the tip of the syringe was damaged. Further description notes that the tip was broken. | NURSE ASSIST INC | 0012555852 | 2017-07-25 | 2017-07-26 | 0 | 2017-07-25 | During use | 2017-10-25 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112985, Serial/Lot No:46132, Description:10 mL water SYR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 10 mL water SYR Model Number: 8881112985 Serial or Lot Number: 46132 ------------------------------------------------------------------------ --------------- Please provide a detailed description of the event the tip of the syringe was damaged. Was the customer notified the product should be returned? Yes Below are the details for Product - Model No:8881112985, Serial/Lot No:45830, Description:10mL water syr IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 10mL water syr Model Number: 8881112985 Serial or Lot Number: 45830 ------------------------------------------------------------------------ -------------- Please provide a detailed description of the event the tip of the syringe was damaged Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: the tip of the syringe was damaged | Rose Matthews | Jeremy Roberts | 42969 | 42970 | Note Creator: Jeremy Roberts Note Date: 8/22/17 Note Description: Addition of Product Analysis. Product Analysis: Two syringes were returned for evaluation (sample 1 and sample 2) along with a non-medtronic catheter. Neither syringe is in originally supplied packaging. 1: Sample 1 was confirmed to be from Mold 58020 cavity 70. The sample was visually inspected for defects. The luer tip is intact with no abnormalities observed and was confirmed to meet ISO 594 luer taper requirements (Gage ID: SR-734, Calibration due 5/16/18). The syringe barrel shoulder was observed to be concaved forcing the rubber tip and plunger rod away from the fully seated position. This deformation is commonly seen during barrel molding and is due to material build-up on mold core pins which causes the syringe barrel to stick during mold ejection. | Complete | Yes
2017-07-26 | 0701980459-10 | 8970Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8970Y evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 26, 2016 to Aug 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and 1 other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids packed prior to sealing a case. The case quantity for 8970Y is 20 eaches/case. The complaint is for 100 eaches. As per customer complaint attachment mPXR-443975-2017-07-26 it stats that product return status as still in use. The product should not be used without lids. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids and/or containers should not be used. | 0701980459 | 0 | Y | Canada | Closed | According to the reporter, on June 1, 2017, the device was missing a hundred lids prior to use. There was no patient involvement. | 0 | 0 | 2017-06-01 | 2017-07-27 | 0 | 2017-06-01 | Prior to Use | 2017-08-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Amin, Krishna PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: We are missing 100 lids of 8970Y for Lifelabs. I spoke to Krishna in your c/s about this issue back on June 1st and she was going to get these lids in, but I still haven’t seen them. Below are the details for Product - Model No:8970Y, Description:8970Y YEL 2 GAL CONT ROTOR LID X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8970Y YEL 2 GAL CONT ROTOR LID X20 Model Number: 8970Y Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No SALESREP USER ID: molind1 SALESREP EMP ID: 252255 SALESREP NAME: Molino, Dayna PRIMARY CONTACT: Someone else PRIMARY CONTACT OTHER: Distributor PRIMARY CONTACT OTHER TEXT: Colleen Hewlin Buyer Tel: 905-791-8600 Ext. 1272 colleen.hewlin@stevens.ca COMPLAINT OWNER NAME: Molino, Dayna As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-26 | 0701980521-10 | 8998 | 00176283 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998 RED 18 GAL CONT HINGE/SEAL LIDX5 | MITG - Sharps Disposal | 002927 | 8998 |  |  |  | RFR | LID ISSUE | AFC | KC:CRACKED LIDS | The Product 8998 evaluated for the cracked lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176283 (05/05/17 to 05/07/17) and no internal reject found during production. Trending Period: July 26, 2016 to Oct 03, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: This issue has been determined to be an isolated event with minimal risk as cracked lids should not be used. Before use container/lid needs to be inspected for any damage per IFU. The potential root cause has been determined to be damage during shipping and handling due to external stress/forces. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as cracked lid product should not be used. The issue was identified prior to use and there is no harm or injury reported to patient safety. | 0701980521 | 0 | Y | United States | Closed | The customer states Lid torn by handle. | Ridgeview Medical Center | 0015922839 | 2017-07-26 | 2017-07-26 | 0 | 2017-07-26 | Unknown | 2017-11-11 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-27 | 0701982959-10 | 8933 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8933 RED 12 GAL CONT W/HINGED LID X10 | MITG - Sharps Disposal | 002927 | 8933 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8933 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 26, 2016 to Aug 16, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8933 is 10 eaches/case. There is no information in the complaint regarding if the customer received the product in original Cardinal Health overshipper case or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701982959 | 0 | Y | United States | Closed | According to the reporter, on July 26, 2017, the packaging of the product had no lids for the containers prior to usage. | Hurley Medical Center | 0008004247 | 2017-07-27 | 2017-08-04 | 0 | 2017-07-26 | Prior to Use | 2017-08-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Trovato, Christopher PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: NO LIDS FOR THE CONTAINERS WERE IN THE PACKAGING Below are the details for Product - Model No:8933, Description:8933 RED 12 GAL CONT W/HINGED LID X10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8933 RED 12 GAL CONT W/HINGED LID X10 Model Number: 8933 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No SALESREP USER ID: trovac1 SALESREP EMP ID: 229312 SALESREP NAME: Trovato, Christopher PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Trovato, Christopher | Justin Moss | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-31 | 0701987990-10 | 8541SA | 00175387 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 4 GA SHARPSTAR IV TRANS RED | MITG - Sharps Disposal | 002927 | 8541SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8541SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175387 (01/30/17 to 02/02/17) and no internal reject found during production for the same issue. Trending Period: July 31, 2016 to Aug 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8541SA is 10 eaches/case. The DHR was reviewed and met acceptance criteria. There is no information in the complaint regarding if the customer received the product in original Cardinal Health overshipper case or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701987990 | 0 | Y | United States | Closed | The customer reports six missing lids. | SURGICAL CENTER OF GREENSBORO | 0013220739 | 2017-07-13 | 2017-07-31 | 0 | 0 | Prior to Use | 2017-08-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-07-31 | 0701988672-10 | 8506Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8506Y YEL 5 QRT CONT S STAR C B DR X20 | MITG - Sharps Disposal | 002927 | 8506Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MALFUNCTIONING LID | The Product 8506Y evaluated for the failure to drop sharp-malfunction lid. The sharps was stuck vertically on one side of door. As from photo evaluation sharps were not disposed horizontally. The reported condition cannot be observed for door function due to no physical sample available for evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 31, 2016 to Aug 16, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The photo was provided in the complaint. The sharps was stuck vertically on one side of door. As from photo evaluation sharps were not disposed horizontally. The cause is customer may not be following Instruction for Use (IFU) to drop sharps horizontally and Lift to assure for disposal. The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. The Instruction for Use (IFU) to drop sharps horizontally and Lift to assure for disposal was not followed by customer. | 0701988672 | 0 | Y | Canada | Closed | According to the reporter, on 31 July 2017, accordingly, the needle was not completely stuck,but would need to jostle and flip the moveable part of the lid to be able to drop the needle in to the container.The lids reading as FULL and staying flipped closed wherein the contents are no where near full line. Ampules got caught in the lid. There was no patient involvement. | Surrey Memorial Hospital | 0008050653 | 2017-07-31 | 2017-08-02 | 0 | 2017-07-31 | During use | 2017-11-04 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: The needle shown in the photo (attached to this email as “sharps2”) was not completely stuck (ie- I was eventually able to drop the needle in to the container), but I had to jostle the container in addition to trying to flip the moveable part of the lid to get the needle to go down in to the container. I do not know how the needle became lodged in that position as I discovered the needle the during the first time I had to use the sharps container on my shift (ie- was placed prior to my shift). Other concerns fellow RN’s have regarding these sharps containers include: - the lids reading as “FULL” / staying flipped closed when the contents of the containers are no where near the full line - ampules getting caught in the lid. Below are the details for Product - Model No:8506Y, Description:8506Y YEL 5 QRT CONT S STAR C B DR X20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Asked but unknown PRODUCT EVENT DETAILS: Event Details for 8506Y YEL 5 QRT CONT S STAR C B DR X20 Model Number: 8506Y Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Asked but unknown Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Dun PHYSICIAN FIRST NAME: Jackie PHYSICIAN CLINIC: Surrey Memorial Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Surrey PHYSICIAN STATE CODE: BC PHYSICIAN EMAIL: Jacqueline.Dun@fraserhealth.ca For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Dahlbeck, Michelle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZINTCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH ZFOLCONT 331345 Dahlbeck, Michelle cannot find a BP in GCH | Daniel Tamul | Gatha Kothari | 42963 | 42963 | The photo was provided in the complaint. The sharps was stuck vertically on one side of door. As from photo evaluation sharps were not disposed horizontally. The reported condition cannot be observed for door function due to no physical sample available for evaluation. | Complete | Yes
2017-08-01 | 0701990529-10 | 8950SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8950SA 5 QRT RED ROUND W/ROTOR OPENING | MITG - Sharps Disposal | 002927 | 8950SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8950SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 1, 2016 to Aug 23, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8950SA is 40 eaches/case. There is no information in the complaint regarding if the customer received the product in original Cardinal Health overshipper case or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0701990529 | 0 | Y | United States | Closed | According to the reporter, there were missing lids. There was no patient involvement. | 0 | 0 | 2017-07-28 | 2017-08-04 | 0 | 2017-07-28 | Prior to Use | 2017-08-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Razzi, Elisabeth PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: There are just missing lids, no patient involved, no facility involved. Below are the details for Product - Model No:8950SA, Description:8950SA 5 QRT RED ROUND W/ROTOR OPENING IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8950SA 5 QRT RED ROUND W/ROTOR OPENING Model Number: 8950SA Serial or Lot Number: null ------------------------------------------------------------------------ --------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Was the customer notified the product should be returned? No Please explain This can not be returned because we are only missing 10 lids and just need 10 replacement lids. SALESREP USER ID: razzie1 SALESREP EMP ID: 231544 SALESREP NAME: Razzi, Elisabeth PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Razzi, Elisabeth | Justin Moss | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-08 | 0702005002-10 | 8506Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8506Y YEL 5 QRT CONT S STAR C B DR X20 | MITG - Sharps Disposal | 002927 | 8506Y |  |  |  | RFR | DOOR FAILS TO FUNCTION | AFC | KC:MALFUNCTIONING LID | The Product 8506Y evaluated for the malfunction lid. The reported condition was observed on photo but as no physical sample available door functionality cannot be evaluated. DHR Review was not performed and met acceptance criteria. DHR Review: The lot number was obtained from the sample photograph. Device History Records data reviewed for 176577 (04/04/17 to 04/06/2017) and no internal reject found during production. The door function inspection was performed during manufacturing and met all acceptance criteria. Trending Period: Aug 2, 2016 to Sep 1, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family/lid type. Root Cause: The DHR review was performed based on lot number information gathered from sample photo evaluation. The ampule is very light weight sharps disposal which required door function “LIFT TO ASSURE DISPOSAL” as printed on the door. It may require different style of lid-container should be used in this type of setting based on type of sharps disposal (ampule-light weight). The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. It may be different style lid/container opening unit should be used for customer’s specific end use application. Also as per instruction for use needs to be follow for door function “LIFT TO ASSURE DISPOSAL”. | 0702005002 | 0 | Y | Canada | Closed | The customer states that ampules are getting stuck in a sharps container lid twice. The customer had to flip the lid 2- 3 times to get the ampule pieces to drop in to the container. | Surrey Memorial Hospital | 0001097356 | 2017-08-02 | 2017-08-08 | 0 | 2017-08-02 | During use | 2017-11-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Janice Nevius | Gatha Kothari | 42979 | 42979 | The Sample photo was evaluated from attachment "FW mPXR # 446892 Notify Date 2017-Aug-02". From the sample photo lot number 176577 was obtained. As per photo ampule was stuck between door edge and the container inside. The ampule seems like very light weight. The door needs to lift for disposal as instructed on door “LIFT TO ASSURE DISPOSAL”. As physical sample is not available so door functionality cannot be evaluated. | Complete | Yes
2017-08-10 | 0702007942-10 | 5551777369MT | 17C06863X | COV LP/LLC, Crystal Lake, Mfg | GENERAL PMR |  | 5551777369MT 10ML SYR LL LOR PTD BARREL | MITG - Needles & Syringes | 003208 | 5551777369MT |  |  |  | RFR | PLUNGER ISSUE | AFC | KC:INCORRECT ASSEMBLY | Note Creator: Jeremy Roberts Note Date: 9/5/17 Note Description: Addition of Analysis. Incident Description: The customer reports the Syringe manufacturing machine stops during the automated assembly process. It blocks, because the plunger cannot be inserted into the barrel as usual. The resistance is either to high, or it is just impossible. The inner diameter of the proximal end of the syringe does not meet our specification. Drawing “B” of the enclosed drawings fits our specs. Back in year 2012 (see your complaint no.s KC201202-0403 and KC201210-0709) we saw the same occurrence. After the reason for the complaint was sorted out (two different molds are available at the manufacturing site) we were assured that if we place an order for item 5551777369 MT LL-Barrel, we will get the barrel we require – which is the barrel without the gate step, see mails attached. Because of the dimensions/diameters we have measured in the inner of the barrel i. e.: 15,22mm 15,15mm 15,16mm 15,19mm and the visual inspection. we are concluding that the WRONG molds were used again! Therefore we are asking for urgent replacement of 120,000 barrels, see invoice attach ed Product Analysis: As of 9/5/17 samples were not returned for evaluation to analyze the reported condition. If sample product of this lot is returned, further testing and investigation may be performed. DHR Review: A review of the device history record was completed for lot 17C06863X. 120 cases (102,000 pieces) on 3/26/17 using production line 7. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process, inclusive of running using validated process settings. DHR review confirmed mold 58140-01 was used during production. Root Cause Analysis: The method of root cause analysis deployed within this investigation was to perform a review of changes deployed on 5551777369MT, going back to when product was manufactured by Commerce Texas. Relevant Change Data: - 5551777369MT has always used Barrel PN 3150420. - 3150420 is made to drawing DR10004386. - DR10004386 has always included a ring stop (gate step) with the current revision requiring the diameter to be 0.596” +/- 0.006. - Change control CDP-APCI-0005808 was completed 11/11/14 to transfer manufacturing from Commerce Texas to Crystal Lake Illinois. o As a corrective action to the Customer complaints noted in the incident description, it was identified that Commerce Texas added s pecial instructions to their MRP (BPCS) to provide guidance to their manufacturing team. They did not house Mold information in the Device Master Record (DMR). Furthermore, they failed to modify the part number and drawing to represent customer requirements. o FS10013324 Rev 4 added mold 58042 on 11/11/14 to support the transfer of manufacturing from Commerce Texas to Crystal Lake Illinois. This ensured the DMR included mold information as required that matched the barrel part number and drawing on the DMR. o Mold 58042 had two configurations:  58042 with ring stop/no fill ring (DR10004386 Barrel PN 3150420).  58042-01 without ring stop (DR10012278 Barrel PN 3150429). - FS10013324 Rev 6 was released 1/26/16 to add mold 58140-01 (drawing DR10004386). - CDP-0002260 decommissioned mold 58042 on 11/17/16 due to being unsafe for use. Root Cause: The device master record was not updated as part of the 2012 customer complaint corrective action to reflect customer requirements (Formula sheet/bill of material/barrel drawing). Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal rejects pertaining to this failure mode (Incorrect Assembly) and product family was performed. Zero other incidents of this nature for this product family was received in the trending period from 6.1.16 – 9.5.17. Zero internal rejects of this nature for this product family were received in the trending period from 4.1.17 – 9.5.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time. CDP-0002260 decommissioned mold 58042 on 11/17/16 due to being unsafe for use. At this time a separate mold which manufactures an LK no ring stop component is not available (new mold would be required). Based on demand, it is recommended that the complaint product code be discontinued. Future orders will include a ring stop if placed, in accordance with the active design specification. Furthermore, preventative CAPA2016-010628 was closed on 2/27/17. Drawing and part number changes were made as part of this CAPA to minimize use of alternate designs while creating a clear distinction within the DMR. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Review: RE00035278 Rev. 0 pFMEA for Syringe Assembly, Rotary Assembly Machine (RAM) was reviewed which confirmed inclusion of failure for incorrect assembly necessary and associated controls. | 0702007942 | 0 | Y | Germany | Closed | the customer reports the Syringe manufacturing machine stops during the automated assembly process. It blocks, because the plunger cannot be inserted into the barrel as usual. The resistance is either to high, or it is just impossible. The inner diameter of the proximal end of the syringe does not meet our specification. Drawing “B” of the enclosed drawings fits our specs. Back in year 2012 (see your complaint no.s KC201202-0403 and KC201210-0709) we saw the same occurrence. After the reason for the complaint was sorted out (two different molds are available at the manufacturing site) we were assured that if we place an order for item 5551777369 MT LL-Barrel, we will get the barrel we require – which is the barrel without the gate step, see mails attached. Because of the dimensions/diameters we have measured in the inner of the barrel i. e.: 15,22mm 15,15mm 15,16mm 15,19mm and the visual inspection. we are concluding that the WRONG molds were used again! Therefore we are asking for urgent replacement of 120,000 barrels, see invoice attached | PreOx.RS GmbH | 0001466674 | 2017-06-01 | 2017-08-10 | 0 | 2017-06-01 | Unknown | 2017-10-04 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Justin Moss | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-10 | 0702009158-10 | 8537SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8537SA 12 QRT CONT S/STAR C/B DR T/RED | MITG - Sharps Disposal | 002927 | 8537SA |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:NEEDLE PUNCTURED CONTAINER | The Product 8537SA evaluated for the needle punctured container. The reported condition was observed in the physical sample. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 08, 2016 to Aug 25, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause is customer misuse by dropping sharps in perpendicular position. Also the container was autoclaved. So the sharps disposal content may have been moved during autoclaving and handling process resulted in the needle punctured through container. As per the sample the container was squeezed due to temperature, pressure and vacuum generated due to autoclave process. Which may result in the needle punctured container. The sharps containers comply with the requirement for puncture resistance defined in ISO 23907 first edition 2012-09-01 Sharps injury protection – Requirements and test methods – Sharps containers. This test method defines the puncture resistance for sharps containers at room temperature, which is the normal use environment. Sharps containers which are designed to allow the contents to be decontaminated by autoclave includes vents to allow steam penetration and allow pressure equalization such that the container will remain intact after the autoclave processing, and prevent spillage of the contents. However, at the extreme temperatures required to achieve decontamination, the material of the sharps container will become softer and the resistance to needle penetration is temporarily reduced until the plastic has cooled to room temperature. Needle within the container that are pressed against the container walls may penetrate through the hot plastic while within the autoclave. The fact that the needle punctured the container during the autoclave process is not unexpected. However, the use of an autoclave to decontaminate the contents of a Sharps Container is only appropriate when it is part of a complete system specifically designed for waste treatment. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event due to customer misuse. The sharps were disposed perpendicular into the container instead of horizontal drop. The fact that the needle punctured the container during the autoclave process is not unexpected. However, the use of an autoclave to decontaminate the contents of a Sharps Container is only appropriate when it is part of a complete system specifically designed for waste treatment. | 0702009158 | 0 | Y | United States | Closed | The customer stated needle went through the side, post autoclaving. | Thermo Fisher Scientific | 0015890421 | 2017-08-08 | 2017-08-10 | 0 | 2017-08-03 | During use | 2017-12-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Non-Serious Illness Injury | 0 | Daniel Tamul | Gatha Kothari | 42972 | 42972 | The physical sample was evaluated. The photos were attached in the complaint. The needle punctured through container in the front side (label side) of container in the middle below the label and near the bottom area. Also two other needles punctured through container on the non-label side of container towards bottom area. There were total four needle punctured through container two in front and two in back side of the container. From the physical sample the container was squeezed due to autoclave. The container was not able to stand on flat surface. The wall thickness measured was 0.067” which is well above the specification. The needle penetration testing was performed and it met the acceptance criteria. The average puncture force 5.98 lbf was greater than 3.4 lbf specification. Refer to test results and physical sample evaluation photos in the attachment. | Complete | #N/A
2017-08-11 | 0702011968-10 | 8989 | 17E07063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E07063 (05/31/17 to 06/05/2017) and no internal reject found during production. Trending Period: July 26, 2016 to Aug 28, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and ten other complaints reported for the product family/different lid types. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8989 is 5 eaches/case. The physical sample is not available also there is no information about if customer receive product in original Cardinal Health packaging or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702011968 | 0 | Y | United States | Closed | The customer states missing lids. | FHN Memorial Hospital | 0015734440 | 2017-07-26 | 2017-08-11 | 0 | 0 | Unknown | 2017-10-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-15 | 0702017524-10 | 8881112599 | 17D04863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING COMPONENTS | Note Creator: Jeremy Roberts Note Date: 9/12/17 Note Description: Addition of Investigation. Incident Description: According to the reporter, six cases of product were missing the rubber bands were open which compromised cleanliness and therefore were rejected. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size for part 8881112599 is 624 cases (499,200 pieces) and production line 1 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The assessment eval uated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Not Significant – Operators are trained to performed required functions in accordance with packaging instructions provided in the design specification). Measurement Causes: Visual Acceptance Criteria (Significant – Criteria exists and is required to be used during manufacturing. It is anticipated that c ustomers align with criteria used during manufacturing as product is supplied in accordance with manufacturing design specifications. The AQL assigned with this defect is 1.5. In the event random sampling is conducted under normal LII sampling plans the customer would acc on 3/rej on 4. With the lot being reported as unknown, it is concluded that the missing rubber bands were identified during 100% inspection prior to loading into their filling operation. With 100% inspection being performed an equivalent acc of 15/rej of 16 should be deployed. AQL has been met for the reported condition). Material Causes: - Rubber Bands (Not Significant – The customer confirmed rubber bands were missing from 6 cases. As such it can be deduced that all other cases supplied (618) included them. The AQL assigned to this failure mode is 1.5. Acceptance criteria has been met. Method Manual Packaging Process (Not Significant – The customer confirmed rubber bands were missing from 6 cases. It can be deduced that all other cases supplied (618) included them. The AQL assigned to this failure mode is 1.5. Acceptance criteria has been met. Further processing at customer location (Significant – The customer confirmed rubber bands were missing from 6 cases. It can be deduced that all other cases supplied (618) included them. It is possible the rubber bands broke while handling at the customer location. Visits of the customer location have been conducted. The customer removes bags from the supplied over shipper cartons and transports them to their further processing environment). Machine Causes: N/A: The related packaging process is completed manually. Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: In the event of a failure during the manufacturing process, the most probable root cause is human error. It is not clear from the information provided if the rubber band broken while handling at the customer’s site, however based on the information it was concluded that random sampling was not performed. It is recommended the customer align to an AQL of 1.5 for this failure. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Missing Components) and product family was performed. One other incidents of this nature for this product family was received in the trending period from 6.29.16 – 9.12.17. Zero internal rejects of this nature for this product family were received in the trending period from 4.29.17 – 9.12.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Review: RE00060107 Rev. 0 pFMEA for Manual Packaging was reviewed. It was confirmed that missing component was not included. Change order RC113155 has been initiated to ensure missing components is covered along with the associated controls. | 0702017524 | 0 | Y | United States | Closed | According to the reporter, six cases of product were missing the rubber bands were open which compromised cleanliness and therefore were rejected. There is no patient involved in this event | AQUABILITI | 0012728353 | 2017-07-05 | 2017-08-16 | 0 | 2017-06-29 | Unknown | 2017-09-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | No Injury | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Quality Engineer Mitchell Eichas ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Cases still at Aquabiliti CONTACT NAME AND PHONE: Mitchell Eichas, ph# 615-833-2633 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event 6 cases of lot# 17D04863X were missing rubber bands and were open which compromised cleanliness and therefore were rejected. Aquabiliti requests a return to Kenmex plant. Was the customer notified the product should be returned? Yes SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Aquabilliti requires a written close-out letter and requests the 6 cases be returned to KM. Please coordinate return with Customer Service. | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-16 | 0702019917-10 | 31143699 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1 GALLON RIGID PLASTIC CONTAINER (RED) | MITG - Sharps Disposal | 002927 | 31143699 |  |  |  | RFR | CONTAMINATION | AFC | KC:CONTAMINATION | The Product 31143699 evaluated for the contamination. The reported condition was observed on the photo having contamination on part. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 14, 2016 to Sept 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error to inspect product for contamination issues prior to seal for the finished good case. The root cause location might be injection molding machine, mold or conveyor had contaminant and it adhered to part during production. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as contaminated product should not be used. The product should be inspected prior to use for any abnormality. | 0702019917 | 0 | Y | United States | Closed | The customer states that item is dirty/ contaminated. | Pacific Biolabs Inc | 0016455849 | 2017-08-14 | 2017-08-16 | 0 | 2017-08-14 | Prior to Use | 2017-10-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | Gatha Kothari | 42986 | 42986 | The photo was submitted with this complaint. The reported condition stated by the customer was observed in the photo. It seems like grease on the lids but it cannot be confirm what type of contamination without actual physical sample. | Complete | #N/A
2017-08-16 | 0702020229-10 | 8820- | 00172173 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 2GAL.PHARMASAFETY WAST CONT.20 | MITG - Sharps Disposal | 002927 | 8820- |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8820 evaluated for the missing pads. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 172173 (08/29/16 to 08/31/16). There was no internal reject found during production. Trending Period: Aug 15, 2016 to Sept 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and 4 other complaints reported for the product family/lid type. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. Also outside of Cardinal Health case may have been broken down to small quantity to ship to end user can cause missing absorbent pad issue. The absorbent pads are 4 packet of 5 each per packet (20 per case) to have one absorbent pad for each container as case quantity is 20 per case. There was no information in the complaint regarding quantity ordered by customer. The defective quantity listed as 2 each which means the customer did not receive full case of 20 eaches and case was broken down outside of Cardinal Health may resulted in missing absorbent pads. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads. There is no patient involvement, so there is no risk to patient safety or harm. | 0702020229 | 0 | Y | United States | Closed | The customer states that they did not receive the absorbent pads in the containers. | Associated Dental | 0016456105 | 2017-08-15 | 2017-08-16 | 0 | 2017-06-07 | Unknown | 2017-10-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-16 | 0702020983-10 | 8938 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8938 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 16, 2016 to Sep 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eleven other complaints reported for the product family/different lid types. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8938 is 5 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702020983 | 0 | Y | United States | Closed | According the reporter, prior to use the lids of the product were missing upon receipt. There is no patient involved in this event | Medical City Weatherford | 0004060071 | 2017-08-16 | 2017-08-17 | 0 | 0 | During use | 2017-09-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | No Injury | FIRST NOTIFIED NAME: Anastasio, Eric PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8938, Description:8938 RED 18 GAL CONT W SLIDING LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8938 RED 18 GAL CONT W SLIDING LID X5 Model Number: 8938 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing upon reciept Was the customer notified the product should be returned? Yes For Physician with Siebel ID 0014717773: PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: anaste1 SALESREP EMP ID: 313874 SALESREP NAME: Anastasio, Eric PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Anastasio, Eric ZINIREP 0014717773 Sondra Wallace cannot find a BP in GCH | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-21 | 0702028369-10 | 8900SA | 00175939 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175939 (4/04/17 to 4/14/17) and no internal reject found during production. Trending Period: Aug 21, 2016 to Sep 13, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and three other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8900SA is 100 eaches/case. The lid box is packaged in with containers placed in the finished good case. Also there is a weight scale for check. The DHR reviewed and all inspection met acceptance criteria. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702028369 | 0 | Y | United States | Closed | According to the reporter, before use, the device's container lids were missing from the case. There was no patient involvement. | 0 | 0 | 2017-08-21 | 2017-08-23 | 0 | 0 | Prior to Use | 2017-09-22 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8900SA, Serial/Lot No:00175939, Description:8900SA 1QT RED CONTAINER PHLEBOTOMYX100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 Model Number: 8900SA Serial or Lot Number: 00175939 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Container lids were missing from the case PHYSICIAN LAST NAME: Napier PHYSICIAN FIRST NAME: Joanna PHYSICIAN CLINIC: Whitesburgh ARH PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 240 Hospital Drive PHYSICIAN CITY: Whitesburgh PHYSICIAN STATE CODE: KY PHYSICIAN POSTAL CODE: 41858 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: Whitesburgh ARH FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 240 Hospital Drive FACILITY CITY: Whitesburgh FACILITY STATE CODE: KY FACILITY POSTAL CODE: 41858 FACILITY DAY PHONE: 6066333500 SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura ADDITIONAL COMMENTS ABOUT THE EVENT: Request for replacement lids | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-21 | 0702028598-10 | 8881112059 | 17F30463X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | FLASH ISSUE | AFC | KC:POOR PRINT | Note Creator: Jeremy Roberts Note Date: 9/29/17 Note Description: Addition of Investigation Summary. Incident Description: Customer reports I wanted to make you aware of a flash and print issue on the subject lot. Please see the attached photo. We have segregated about 150 pieces at point of use. Product Analysis: As of 9/29/17 samples were not returned for evaluation. Pictures provided with the complaint were sufficient for evaluation against Quality Inspection Standards. Print: In accordance with GS10000795 Revision 13, Inspection Standard for Printing of Syringe Barrel, the pictures provided demonstrate that the printing is within specification and would be classified as a process indicator (scratches on print). Flash: In accordance with GS10005108 Revision 8, Quality and Inspection Standard for Plastic Molded Parts, the pictures provided demonstrate that the flash is within specification of 1/16” in a non-functional area. DHR Review: A review of the device history record was completed for lot 17F30463X. The lot released 288 cases (144,000 pieces) on 7/18/17 using production line 7. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: Reference product analysis. A defect could not be determined at this time. Nothing was discovered at the time of inspection or during the production run that would lead to the formulation of a root cause as product would not be released unless acceptance criteria is met. Poor Print and Flash are common to Syringe manufacturing and evaluated throughout the lot to ensure specifications and acceptable quality levels are attained. With product being segregated at point of use as reported, this is an indication that 100% inspection is being performed. 100% inspection is not feasible in manufacturing. Root Cause: Root cause investigation could not confirm a manufacturing defect as the pictures provided confirmed acceptance criteria was met. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Poor Print and Flash) and product family was performed. Complaints: Four incidents of Poor Print (low severity) and one incident of Flash (low severity) was noted for this product family in the trending period from 8.11.16 – 9.29.17. NCR’s: Zero incidents of Printing/Flash were noted for this product family in the trending period from 7.11.17 – 9.29.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, addition actions are deemed unnecessary at this time as the reported condition for this product family presents a low risk level. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed inclusion of failure for flash and associated controls. RE00035287 Rev. 0 pFMEA for Syringe Pad Printing was reviewed which confirmed inclusion of poor print and associated controls. | 0702028598 | 0 | Y | United States | Closed | the customer states that there's a flash issue with the subject lot. | MERIT MEDICAL SYSTEMS INC | 0013142362 | 2017-08-11 | 2017-08-21 | 0 | 0 | Unknown | 2018-01-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Paul Deangelis | Charlean Favela | 43082 | 43082 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2017-08-23 | 0702033455-10 | 31143665 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31143665 10 GAL RED CONTAINER SHARPSX6 | MITG - Sharps Disposal | 002927 | 31143665 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31143665 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: July 24, 2016 to Sept 06, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 31143665 is 6 eaches/case. There is no information about if the customer received product in Cardinal Health original overshipper case or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702033455 | 0 | Y | United States | Closed | The customer states missing lids. | HARBOR HEIGHTS SURGERY CENTER | 0001479773 | 2017-07-24 | 2017-08-23 | 0 | 0 | Unknown | 2017-09-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Paul Deangelis | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-24 | 0702036275-10 | 851201 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 851201 CLEAR 5 QUART CONT OPEN LID X20 | MITG - Sharps Disposal | 002927 | 851201 |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | MALFUNCTIONING LID | The Product 851201 evaluated for the lid fit issue. The reported condition cannot be observed as no photo or sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 24, 2016 to Sep 22, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: Unable to determine the root cause as the information provided in the complaint is insufficient and also there is no photo or sample available for the evaluation. As from the attachment customer states about “Some of the lids to the sharps containers are defective and do not close properly. The closing mechanism is hard to explain but upon inspection found 2 more lids that are not fully put together. When you try to apply pressure to seal the container a gap is formed right where the lid meets the top of the container.” It is not very clear if the issue is with lid fit with the container or with final lock mechanism. Also as per IFU lid and container needs to be inspected for any shipping damage prior to use. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as there was no patient involvement or consequence to the patient. As no sample or photo is available for the product analysis unable to determine the cause. | 0702036275 | 0 | Y | United States | Closed | The customer states that 3 of the lids are defective in that they do not close properly. When you try to apply pressure to seal the container, a gap is formed right where the lid meets the top of the container. | OSF Multi Specialty Poplar Grove | 0016436918 | 2017-08-24 | 2017-08-24 | 0 | 2017-08-24 | Unknown | 2018-02-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Daniel Tamul | Gatha Kothari | 43110 | 43110 | The 2 lid assembly were received. The lid fit was performed as intended. While performing final lock gap was observed. The three post on the lid were not welded causing gap during final lock. The reported condition was observed. | Complete | Yes
2017-08-24 | 0702036684-10 | 8985CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8985CRS 8 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8985CRS |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:NO FAILURE | The Product 8985CRS evaluated for the no failure. The reported condition cannot be observed from photo provided. The photo provided the final lock was performed as intended product use. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 23, 2016 to Sep 20, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family/different type lids. Root Cause: The probable root cause is IFU and instruction label for the final lock states “to use center of circular door to turn counterclockwise until door drops sown then press firmly in center of door to engage final lock. It is slightly difficult to hold circular door from the center. If the instruction have word “tab” instead of “center of circular door” it may be ease of use for final user. But the instruction were sufficient to perform final lock as per product intended use. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. The marketing will look into any future improvement for instructions. Conclusion: The complaint has RFR listed as “Lid Does not attach/Stay attached” which is incorrect as lid fit was already performed. This issue has been determined to be an isolated event with minimal risk as there was no patient involvement. Also as per the sample photo evaluation final lock was completed as per intended product use. | 0702036684 | 0 | Y | Canada | Closed | According to the reporter, on August 23, 2017, during use, the device had unclear instructions on how to properly lock and dispose of container once it was full. There should be improvements that can be made to the instructions under the “to lock for final disposal”. The first step should be to turn the clear tab to the left as far as it will go. The second step was to use the centre. Third would be to press. Fourth, Lid should be locked to properly so that the dots are clearly visible through the lid. If it was not done, and the clear tab did not go to the left, the lid didn’t lock down properly and you didn’t see the dots clearly through the lid. This had happened on numerous occasions, and could not be corrected. Sometimes, when centre of the circular door was turned, the clear tab did not turn with it to the left, but not consistently. In addition, the reporter was looking forward that this feedback was useful and will look forward for the team to review the product and would agree with the observations. There was no patient involvement. I | Burnaby Hospital | 0001096672 | 2017-08-23 | 2017-08-29 | 0 | 2017-08-23 | During use | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: I do feel that there are improvements that can be made to the instructions under the “To Lock for Final Disposal”. The first step should be to turn the clear tab to the left as far as it will go. The second step to use the centre…. The third step to press… Forth, lid is locked when…. This ensures that the lid locks properly so that the dots are clearly visible through the lid. If you don’t do this, and the clear tab doesn’t go to the left, the lid doesn’t lock down properly and you don’t see the dots clearly through the lid. This had happened on numerous occasions, and can’t be corrected. Sometimes when you turn the centre of the circular door……the clear tab does turn with it to the left, but not consistently. I hope this feedback is useful to you. I look forward to your team reviewing the product and seeing if they agree with my observations. Below are the details for Product - Model No:8985CRS, Description:8985CRS 8 GAL CHEMO RED SLIDE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: Asked but unknown PRODUCT EVENT DETAILS: Event Details for 8985CRS 8 GAL CHEMO RED SLIDE LID Model Number: 8985CRS Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Unclear instructions on how to properly lock and dispose of container once it is full. PHYSICIAN LAST NAME: unrau PHYSICIAN FIRST NAME: Kay PHYSICIAN CLINIC: Burnaby Hospital Cancer Centre PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Burnaby PHYSICIAN STATE CODE: BC PHYSICIAN EMAIL: kay.unrau@fraserhealth.ca For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Dahlbeck, Michelle ADDITIONAL COMMENTS ABOUT THE EVENT: Please send final report to rep Michelle Dahlbeck michelle.a.dahlbeck@medtronic.com when investigation has been completed. Thank you. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | Gatha Kothari | 42998 | 42998 | The photos were provided with the complaint. The lid final lock was performed per photo 107490-clear dots. The dots are visible clear so there is no form, fit and function issue. The IFU and instruction on the slide top lid assembly was reviewed. The instructions are adequate to perform final lock. Some improvement can be made in future. But per complaint photos sample evaluation there was no defect found with product for the intended use. | Complete | Yes
2017-08-28 | 0702041810-10 | 8990Y | 8990Y | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8990Y YEL 2 GAL BIOHAZ W/O GASKET X20 | MITG - Sharps Disposal | 002927 | 8990Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8990Y evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 28, 2016 to Sep 15, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and nine other complaints reported for the product family/different type lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8990Y is 20 eaches/case. The DHR cannot be reviewed as lot number was unknown. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702041810 | 0 | Y | Canada | Closed | According to the reporter, the sharps container lids were not received with shipment. It was missing prior to use. There was no patient involvement. | PUBLIC HEALTH SERVICES | 0013205453 | 2017-08-28 | 2017-08-29 | 0 | 0 | Prior to Use | 2017-09-26 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Russell-Mondor, Joanne EVENT DATE: Unknown, but will follow up for more information Event Context: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: through customer service via the customer ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8990Y, Serial/Lot No:8990Y, Description:Sharps container 2 Gal IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for Sharps container 2 Gal Model Number: 8990Y Serial or Lot Number: 8990Y ------------------------------------------------------------------------ ----------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event sharps container lids were not received with shipment. Missing prior to use Was the customer notified the product should be returned? Yes SALESREP USER ID: russej2 SALESREP EMP ID: 331012 SALESREP NAME: Russell-Mondor, Joanne PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Russell-Mondor, Joanne ADDITIONAL COMMENTS ABOUT THE EVENT: missing sharps lids from order As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-31 | 0702049288-10 | 8881512878 | 16E07863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | PACKAGING | AFC | KC:MISCOUNT | Note Creator: Jeremy Roberts Note Date: 10/16/17 Note Description: Addition of Investigation Summary. Incident Description: Customer reports we discovered 1 box of 12 ml syringes 8881512878 that felt not as full as the others. I opened it from the bottom and there were only 50 syringes in an 80 count box. Product Analysis: As of 10/16/17 a sample was not returned for analysis. Two pictures were provided with the complaint. The pictures show a sealed inner carton of 8881512878 along with missing product when viewed from the opened bottom as reported by the customer. DHR Review: A review of the device history record was completed for lot 16E07863X. The lot released 326 cases (156,480 pieces) on 7/18/16 using production line 23. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes. Significant causes were then noted and addressed in the corrective and preventative action section as appropriate. Man Causes: - Training - Improper setting of sensors or air (Not Significant – improper setting would be easily detected through line/machine outputs) - Clearing of auto packer jams/picker faults (Significant – In the event of a jam/picker fault, manual intervention is required to remove the inner case for verification of conformance.) Material Causes: - RAM sending damaged parts to packer (Not Significant – Jam will result in machine stop, 2x Sampling Inspection conducted for process verification.) - Damaged Inner cartons (Not Significant – Cartons go through incoming inspection and are only released if quality standards are met.) Method Causes: - Preventative/Maintenance – RAM/Packer (Not Significant – Preventative maintenance is in place to insure equipment is operating properly). - Continued operation – RAM/Packer (Not Significant – Process specification and standard work is in place. Standard work was confir med to be posted at point of use.) Measurement Causes: - Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). - Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot). Machine Causes: - Speeds, feeds, air, vacuum (Not significant – speeds, feeds, air and vacuum are kept within process control ranges where applicable. Other changes would be detected through line outputs). - Packer Jams: (Not Significant – Jams are part of common cause variation. Processes are in place to address jams. - RAM Sending unassembled parts to Packer (Not significant – The RAM is designed to catch unassembled parts. Redundant (2x) sensors are in place on Packer to ensure correct quantity, Machine will fault if missing sleeve). Environment Causes: -Humidity and Temperature – (Not significant – humidity and temperature are controlled in the CME). Root Cause: The most probable root cause is insufficient clearing of auto packer jams/picker faults due to human error. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Miscount) and product family was performed. One other incident of this nature for this product family was received in the trending period from 8.30.16 – 10.16.17. Zero internal rejects of this nature for this product family were received in the trending period from 5.30.17 – 10.16.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time as the reported condition presents a low risk level. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035282 Rev. 0 pFMEA for Syringe Auto Packing was reviewed which confirmed inclusion of failure for incorrect quantity and associated controls. | 0702049288 | 0 | Y | United States | Closed | The customer states that there were only 50 syringes in an 80 count box. | MIDWEST VETERINARY SUPPLY | 0013104364 | 2017-08-30 | 2017-08-31 | 0 | 2017-08-30 | Unknown | 2017-11-11 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-08-31 | 0702050009-10 | 8881112059 | 15G09463X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | FLUID PATH CONTAMINATION | AFC | KC:COSMETIC ISSUE | Note Creator: Jeremy Roberts Note Date: 10/16/17 Note Description: Addition of Investigation Summary. Incident Description: According to the reporter, on August 31, 2017, during use, the device had black spots on syringe. There was no patient involvement. The customer further states that the spots are embedded in the barrel. Product Analysis: One syringe was returned for evaluation. The syringe was not in originally supplied packaging. Upon visual inspection of the syringe subassembly, the plunger rod was observed to have molded in inclusions. Being on the plunger rod, the molded inclusion is classified to be in a non-critical and non-fluid path area. These inclusions are considered as process indicators, opposed to defects. The plunger sample was confirmed to be cavity from 29B. DHR Review: A review of the device history record was completed for lot 15G09463X. The lot released 180 cases (90,000 pieces) on 8/2/15 using production line 7. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: Based on the product analysis, a defect could not be determined at this time. Molded inclusions in non-critical areas such as the plunger rod are considered a process indicator, but acceptable. Process indicators are non-defect conditions that identify a characteristic which does not affect the form, fit or function of a product. This characteristic indicates a condition that, while not necessarily perfect, will maintain the integrity and reliability of the product in its use environment. Molded parts are inspected for contamination/specs throughout the lot to ensure specifications are met. Root Cause: Molded inclusions occur after tool maintenance activities and during molding startup activities. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (cosmetic issue/contamination) and product family was performed. Zero other incident of this nature for this product family was received in the trending period from 8.31.16 – 10.16.17. Zero internal rejects of this nature for this product family were received in the trending period from 5.31.17 – 10.16.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time as the reported condition presents a low risk level. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035275 Rev. 0 pFMEA for Syringe Molding was reviewed which confirmed inclusion of failure for contamination and associated controls. | 0702050009 | 0 | Y | United States | Closed | According to the reporter, on August 31, 2017, during use, the device had black spots on syringe. There was no patient involvement. The customer further states that the spots are embedded in the barrel. | BARD SHANNON LIMITED | 0012818702 | 2017-08-31 | 2017-09-01 | 0 | 2017-08-31 | During use | 2017-10-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | No Injury | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Serial/Lot No:15G09463X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 15G09463X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event BLACK SPOTS ON SYRINGE Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas ADDITIONAL COMMENTS ABOUT THE EVENT: BLACK SPOTS ON SYRINGE | Daniel Tamul | Jeremy Roberts | 43024 | 43024 | Note Creator: Jeremy Roberts Note Date: 10/16/17 Note Description: Addition of analysis. Product Analysis: One syringe was returned for evaluation. The syringe was not in originally supplied packaging. Upon visual inspection of the syringe subassembly, the plunger rod was observed to have molded in inclusions. Being on the plunger rod, the molded inclusion is classified to be in a non-critical and non-fluid path area. These inclusions are considered as process indicators, opposed to defects. The plunger sample was confirmed to be cavity from 29B. | Complete | #N/A
2017-09-01 | 0702051583-10 | 8939 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8939 18GAL WHI CONT CHEMO SLIDE TOPX5 | MITG - Chemo Protection Safety | 003209 | 8939 |  |  |  | RFR | LID ISSUE | AFC | KC:NO FAILURE | The Product 8939 evaluated for the no failure. The reported condition was not observed in the photo as gasket specification are changed per CDP-APCI-0005007 and per photo it seems per current criteria. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Aug 28, 2016 to Sep 15, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: There is no issue found with the gasket appearance. The gasket specification was changed per CDP-APCI-0005007 and all required valid ation and testing met acceptance criteria. No fault found. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as gasket is per current specification. No fault found in this customer complaint. | 0702051583 | 0 | Y | Japan | Closed | The customer reported The rubber inside the rid was thinner than expected. prior to use. No patient involvement | Tsuchiura Kyodo General Hospital | 0008028674 | 2017-09-01 | 2017-09-01 | 0 | 2017-08-31 | Prior to Use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | Gatha Kothari | 42993 | 42993 | The sample photo was evaluated. The gasket specification are changed per CDP-APCI-0005007. From the photo the gasket appearance per current specification. As no physical sample available for gasket dimension evaluation. If sample is available the product analysis summary will be updated. | Complete | #N/A
2017-09-01 | 0702052321-10 | 8906 | 00176515 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 8906 YEL 1 QRT CONT PHLEBOTOMY X100 | MITG - Sharps Disposal | 002927 | 8906 |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | KC:MALFUNCTIONING LID | The Product 8906 evaluated for the malfunction lid. The reported condition cannot be observed from photo provided. DHR Review was pe rformed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176515 (04/14/17 to 04/21/17) and no internal reject found during production. Trending Period: Aug 31, 2016 to Sep 28, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The probable root cause is user did not follow temporary/final lock and carrying/lifting instructions per instruction for use (IFU). Also there is not much information available for product analysis. The information in the complaint is minimal to perform investigation for the failure without sample. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. As per complaint there was no patient involvement. | 0702052321 | 0 | Y | Canada | Closed | According to the reporter, on August 31, 2017, during use, the container did not close optimally. If home care nurses brought it in the car, the container could open in the trunk. The system unscrewed the butterfly to barrel and did not work well. The risks for an accident were high. There was no patient involvement. | CLSC CHSLD RIMOUSKI-NEIGETTE | 0012958639 | 2017-08-31 | 2017-09-05 | 0 | 2017-08-31 | During use | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Daigle, Billy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8906, Serial/Lot No:00176515, Description:8906 YEL 1 QRT CONT PHLEBOTOMY X100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8906 YEL 1 QRT CONT PHLEBOTOMY X100 Model Number: 8906 Serial or Lot Number: 00176515 ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event This container doesn't close optimaly. If home care nurses bring it in the car the container can open in the trunk. The system to unscrew the butterfly to barrel does not work well. The risks for an accident are high. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: St-Pierre PHYSICIAN FIRST NAME: katherine PHYSICIAN CLINIC: Cisss du bas-St-Laurent PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: rimouski PHYSICIAN STATE CODE: QC PHYSICIAN DAY PHONE: 418-724-5231 Poste 557 PHYSICIAN EMAIL: katherine.st-pierre.cisssbsl@ssss.gouv.qc.ca For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes SALESREP USER ID: daiglb1 SALESREP EMP ID: 322696 SALESREP NAME: Daigle, Billy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daigle, Billy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | Gatha Kothari | 43006 | 43006 | The photos was provided but it is not for the sharp container product. There is no information can be obtained from the photo. The photo is showing just human hand. No product photo is available for product analysis. | Complete | Yes
2017-09-05 | 0702055725-10 | 31323176 | 17E39463 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | SAG14-827-122 FISHER 5QT RED | MITG - Sharps Disposal | 002927 | 31323176 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:NO FAILURE | The Product 31323176 evaluated for the no failure. The reported condition was not observed in the photo as product was packaged per the standard work for packaging FM20008553. DHR Review was not performed as lot number provided was not for the product code 31323176. DHR Review: A Device History Record was not reviewed as lot number 17E39463 was provided which does not belong to product code 31323176, 5 QT. Fisher Container Red Tortous Lid. Trending Period: Sep 5, 2016 to Sep 20, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: There is no issue found with the product packaging. The Cardinal Health sold the product as case not by eaches. The product was packaged as per the standard work for packaging FM20008553 for 31323176 case of 30 eaches. The potential root cause is user is breaking down product case into eaches for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as product was packaged as per the standard work for packaging FM20008553 for 31323176 case of 30 eaches. This is a user issue. | 0702055725 | 0 | Y | United States | Closed | The customer states that the product sold as each - comes in a case of 30 each - product needs a top and a base. | Thermo Fisher Scientific | 0015890421 | 2017-09-05 | 2017-09-05 | 0 | 0 | Unknown | 2017-09-29 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | Gatha Kothari | 42998 | 42998 | The sample photos were evaluated. The 31323176 was packaged as per packaging standard work FM20008553. There was no failure observed from the photo. From the complaint attachment “FW Complaint # 005419M RESPONSE REQUIRE” it was listed as follows: “Allegation: Our warehouse reported product sold as each - comes in a case of 30 each - product needs a top and a base - when you open the case all of the tops are stacked - you have to pull all of them out to get to the bases - this causes extra work in the facility and needs more space to stage - also high risk of errors because you have to dig to the bottom of each case to get a red base to make it a complete unit. Pictures are attached” The product was intended to sell as case by Cardinal Health not by eaches. | Complete | #N/A
2017-09-07 | 0702062175-10 | 8935 | 17F14063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8935 12GAL RED CONTAINERNEEDLEDISPX10 | MITG - Sharps Disposal | 002927 | 8935 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8935 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17F14063 (07/31/17 to 08/03/17) and no internal reject found during production for the same issue. Trending Period: Sep 07, 2016 to Sep 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family/different type lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8935 is 10 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702062175 | 0 | Y | United States | Closed | According the reporter, the product was received with missing (5) five lids instead of 10 lids as specified. There is no patient involved in this event. | RUSSELL COUNTY HOSPITAL | 0013237095 | 2017-09-07 | 2017-09-08 | 0 | 0 | Prior to Use | 2017-10-11 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | No Injury | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8935, Serial/Lot No:17F14063, Description:8935 12GAL RED CONTAINERNEEDLEDISPX10 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8935 12GAL RED CONTAINERNEEDLEDISPX10 Model Number: 8935 Serial or Lot Number: 17F14063 ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Customer received a case of 8935 (10 eaches), but only received 5 lids. She would like 5 replacement lids. For Physician with Siebel ID 0014532623: PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura ZINIREP 0014532623 Stella Pennington cannot find a BP in GCH | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-07 | 0702062208-10 | 31142222 | 00174493 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 2 GAL CONTAINER; RED | MITG - Sharps Disposal | 002927 | 31142222 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING COMPONENTS | The Product 31142222 evaluated for the missing component. The reported condition missing door component was observed in the photo. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174493 (11/16/16 to 11/23/2016) and no internal reject found during production. Trending Period: March 2, 2016 to Sept 12, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the door component during slide top lid assembly operation. There was a potential failure during the manufacturing of lid assembly. An associate may have missed an assembly step prior to packaging. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing component product should not be used. | 0702062208 | 0 | Y | United States | Closed | The customer states missing the sliding part of the lid that locks. | HENDRICKS COUNTY HEALTH DEPT | 0013030091 | 2017-03-02 | 2017-09-07 | 0 | 2017-03-02 | Unknown | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Daniel Tamul | Gatha Kothari | 42990 | 42990 | No samples were submitted with this complaint. The photo was attached in the complaint. The lid was missing sliding door component. The defect was observed in the photo. | Complete | Yes
2017-09-08 | 0702064139-10 | 8881412012 | 17F16163X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881412012 CURVED TIP SYRINGE X50 | MITG - Needles & Syringes | 003208 | 8881412012 |  |  |  | RFR | OCCLUDED | AFC | KC:DIFFICULT TO DRAW FLUIDS | Note Creator: Jeremy Roberts Note Date: 10/16/17 Note Description: Addition of Investigation Summary. Incident Description: The customer reported the plunger is not drawing up any medication. Product Analysis: As of 10/16/17 no samples or pictures have been provided. Without samples, we are unable to confirm the reported condition. If samples are received, the complaint may be re-opened and the investigation continued. DHR Review: A review of the device history record was completed for lot 17F16163X. The lot released 240 cases on 6/28/17 using production line 412. The DHR was reviewed, and no abnormal process conditions were present during the manufacturing of this product that would lead to the reported condition. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Fish Bone Diagram that addressed the Six M’s: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). Causes were then categorized and addressed in the corrective and preventative action section where deemed appropriate. Man Causes: Quality and Inspection Standard Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) High Cavitation Injection Mold Processing Training (Not Significant – all applicable associates are trained to perform responsibilities and have training records) Material Causes: Variation is resin viscosity (Not Significant – Resin is evaluated and certified at point of manufacturing) Method Causes: Incorrect/Incomplete Press/Mold Start-up (Not Significant – Records indicate that start-up was completed as expected) Insufficient Preventative Maintenance Process (Not Significant – Preventative maintenance is scheduled and completed through Maximo) Incorrect/Incomplete Mold Repair (Not Significant - Records indicate that start-up was completed as expected) Incorrect Parameter Setting on Mold Setup Sheet (Not Significant – Parameters are verified on a regular to basis to ensure correctness) Insufficient Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Functional/Dimensional Acceptance Criteria (Not Significant – Criteria exists and was used during the production lot) Machine Causes: Damaged Mold Shut Off (Significant – Root Cause) Damaged Mold Core Pin(s) (Significant) Press Challenges Environment Causes: N/A – Production is completed in a controlled manufacturing environment. Formal Investigation: Upon review of the complaint, the evaluation concluded that no further action is required as the below corrective actions are currently in process to address the root cause (assessment completed/solution known). Corrective and Preventive Action: Based on the information available and the investigation findings, a Corrective and Preventive Action (CAPA) is not deemed necessary at this time as the reported condition for this product family presents a low risk level and tool replacement is currently in process under CER 97089 (replacing 412 Tool). If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00035275 Rev. 0. Syringe - Molding (& Setup) was reviewed which confirmed the failure and root cause for “Broken/Damage Non-Functional, Machine Malfunction” is reflected. | 0702064139 | 0 | Y | United States | Closed | The customer reported the plunger is not drawing up any medication. | Dr Mark Flack and Stone | 0016409195 | 2017-09-07 | 2017-09-08 | 0 | 2017-08-01 | During use | 2017-11-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-12 | 0702068571-10 | 8980S | 17E04963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980S 8GAL RED CONT BIOHAZ SLIDETOPX10 | MITG - Sharps Disposal | 002927 | 8980S |  |  |  | RFR | PACKAGING | AFC | KC:DAMAGE PACKAGE | The Product 8980S evaluated for the damaged product. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as invalid lot number provided. DHR Review: Device History Records data cannot be reviewed for 17E04963 as the lot number is invalid for product code 8980S. Trending Period: Sep 11, 2016 to Oct 06, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces. As there is no sample or photos available to perform further investigation. As per the complaint attachment “MCR17J0296／702068571／228312310 Follow up” only outer box was damaged. There was no damage to any of the product inside the box. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk. The product should be inspected prior to use for any damage as per IFU. As per complaint data there was no product damage or no patient involvement with this complaint. | 0702068571 | 0 | Y | Japan | Closed | Customer reports that box was found damaged. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-09-12 | 2017-09-12 | 0 | 2017-09-11 | Prior to Use | 2017-11-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-12 | 0702068574-10 | 8527R | 17E24563 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8527R 12 QRT TRANS RED W/HINGED LIDX10 | MITG - Sharps Disposal | 002927 | 8527R |  |  |  | RFR | PACKAGING | AFC | KC:DAMAGED PRODUCT | The Product 8527R evaluated for the damaged product. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E24563 (07/11/17 to 07/20/17) and no internal reject found during production. Trending Period: Sep 11, 2016 to Sep 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces. As there is no sample or photos available to perform further investigation. As per the complaint attachment “RE MCR17J0298／702068574／228312759 Follo” only outer box was damaged. There was no damage to any of the product inside the box. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk. The product should be inspected prior to use for any damage as per IFU. As per complaint data there was no product damage or no patient involvement with this complaint. | 0702068574 | 0 | Y | Japan | Closed | Customer reports that the box was damaged prior to use. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-09-12 | 2017-09-12 | 0 | 2017-09-11 | Prior to Use | 2017-10-05 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-12 | 0702068999-10 | 8935 | 17E18363 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8935 12GAL RED CONTAINERNEEDLEDISPX10 | MITG - Sharps Disposal | 002927 | 8935 |  |  |  | RFR | PACKAGING | AFC | KC:DAMAGED PRODUCT | The Product 8935 evaluated for the damaged product. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E18363 (06/28/17 to 07/03/17) and no internal reject found during production. Trending Period: Sep 11, 2016 to Sep 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces. As there is no sample or photos available to perform further investigation. As per the complaint attachment “RE MCR17J0298／702068574／228312759 Follo” only outer box was damaged. There was no damage to any of the product inside the box. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk. The product should be inspected prior to use for any damage as per IFU. As per complaint data there was no product damage or no patient involvement with this complaint. | 0702068999 | 0 | Y | Japan | Closed | Customer reports that the box was found damaged prior to use. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-09-12 | 2017-09-12 | 0 | 2017-09-11 | Prior to Use | 2017-10-05 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-13 | 0702072815-10 | 8900SA | 00175939 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175939 (4/04/17 to 4/14/17) and no internal reject found during production. Trending Period: Aug 21, 2016 to Sep 13, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and four other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8900SA is 100 eaches/case. The lid box is packaged in with containers placed in the finished good case. Also there is a weight scale for check. The quantity is 15 eaches per complaint. This shows the Cardinal Health case was broken down for further distribution to the end user by the 3rd party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702072815 | 0 | Y | United States | Closed | The customer states that has advised that they received this container and there were no lids. | St Vincent Hospital | 0015924580 | 2017-09-13 | 2017-09-13 | 0 | 2017-09-12 | Unknown | 2017-10-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-19 | 0702084382-10 | 8900SA | 8900SA | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition stated by the customer was observed during product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175939 (04/04/17 to 04/14/17) and no internal reject found during production for the same issue. Trending Period: Sep 19, 2016 to Oct 9, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8900SA is 100 eaches/case. The lid box is packaged in with containers placed in the finished good case. Also there is a weight scale for check. The probable root cause may be lid box was not placed inside the overshipper box prior to containers. Also there may be missed weight scale check before sealing the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702084382 | 0 | Y | United States | Closed | According to the reporter, on September 19, 2017, pre-operatively, the shipped case had no lids within the container. There was no patient involvement during this event noted. | SENECA MEDICAL INC | 0015731256 | 2017-09-19 | 2017-09-20 | 0 | 2017-09-19 | Prior to Use | 2017-10-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Gill, Chris PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Customer Service Representative ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8900SA, Serial/Lot No:8900SA, Description:8900SA 1QT RED CONTAINER PHLEBOTOMYX100 IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Distributor rep found no lids when putting these on the shelf CONTACT NAME AND PHONE: Sandy Bowles 1-800-445-1044 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 Model Number: 8900SA Serial or Lot Number: 8900SA ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? No When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Shipped case had no lids within the container Was the customer notified the product should be returned? Yes SALESREP USER ID: gillc5 SALESREP EMP ID: 341090 SALESREP NAME: Gill, Chris PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gill, Chris | Shanika Hodge | Gatha Kothari | 43017 | 43017 | The samples were received. The original overshipper case had 100 containers present in the box. The lid box was missing from the finished good overshipper case. Lot number 175939 was obtained from the samples. The reported condition stated by the customer was observed during product analysis. | Complete | #N/A
2017-09-21 | 0702088993-10 | 8980S | 17E04963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8980S 8GAL RED CONT BIOHAZ SLIDETOPX10 | MITG - Sharps Disposal | 002927 | 8980S |  |  |  | RFR | LID ISSUE | AFC | KC:MALFUNCTIONING LID | The Product 8980S evaluated for the malfunction lid (final lock) issue. The reported condition was observed on the sample provided. DHR Review cannot be performed as lot number provided is invalid. DHR Review: Device History Records data cannot be reviewed as lot number provided is invalid for 8980S. Trending Period: Sep 21, 2016 to Oct 11 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product. Root Cause: A probable root cause has been determined to be the lid final lock may happen during shipping and handling of product. It is minimal risk as lid cannot be used and must be replaced. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as final lock lid product should not be used. There was no known impact or consequence to patient as per complaint. | 0702088993 | 0 | Y | Japan | Closed | The customer reports the lid of the container was locked, therefore the container could not be used at all. No patient involved. | 0 | 0 | 2017-09-21 | 2017-09-21 | 0 | 2017-09-14 | Prior to Use | 2017-11-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Janice Nevius | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702095853-10 | V8536 | 175968 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | VHA PLUS SSII 3GA CLR INRM 10 | MITG - Sharps Disposal | 002927 | V8536 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING CONTAINER | The Product V8536 evaluated for the missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175968 (04/18/17 to 04/23/17) and no internal reject found during production for the same issue. Trending Period: Sep 22, 2016 to Oct 9, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of containers during rework process. The affected containers were scrapped but missed to pack new containers with lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing containers product should not be used. | 0702095853 | 0 | Y | United States | Closed | The customer states that they just received the lids and not the container. | MEDLINE INDUSTRIES | 0012965117 | 2017-09-22 | 2017-09-25 | 0 | 0 | Prior to Use | 2017-11-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702096227-10 | 8931PG2 | 00174640 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8931PG2 12GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8931PG2 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8931PG2 evaluated for the missing lids/missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174640 (04/14/17 to 04/18/17) and no internal reject found during production for the same issue. Trending Period: Sep 22, 2016 to Oct 9, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids/containers prior to ship to end user. The case quantity for 8931PG2 is 10 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids/containers product should not be used. | 0702096227 | 0 | Y | United States | Closed | The customer states that factory sealed but only contained 5 each instead of 10 each. | St Mary Medical Ctr | 0012668340 | 2017-09-25 | 2017-09-25 | 0 | 2017-09-22 | Prior to Use | 2017-10-25 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702096227-20 | 8931PG2 | 20884521022918 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8931PG2 12GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8931PG2 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8931PG2 evaluated for the missing lids/missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number provided is invalid. DHR Review: Device History Records data cannot be reviewed as lot number (20884521022918) provided is invalid for 8931PG2. Trending Period: Sep 22, 2016 to Oct 9, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids/containers prior to ship to end user. The case quantity for 8931PG2 is 10 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids/containers product should not be used. | 0702096227 | 0 | Y | United States | Closed | The customer states that factory sealed but only contained 5 each instead of 10 each. | St Mary Medical Ctr | 0012668340 | 2017-09-25 | 2017-09-25 | 0 | 2017-09-22 | Prior to Use | 2017-10-25 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702096699-10 | 8900SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: April 03, 2016 to Oct 03, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user in supply chain. The case quantity for 8900SA is 100 eaches/case. The lid box is packaged in with containers placed in the finished good case. Also there is a weight scale for check. The quantity is 50 eaches per complaint. This shows the Cardinal Health case was broken down for further distribution to the end user by the 3rd party distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702096699 | 0 | Y | United States | Closed | Customer reports that the containers came with no lids. He requests the lids to be sent. | 0 | 0 | 2017-04-03 | 2017-09-28 | 0 | 2017-04-03 | Prior to Use | 2017-10-12 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702097123-10 | 8989PG2 | 175963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID | MITG - Sharps Disposal | 002927 | 8989PG2 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8989PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175963 (02/09/17 to 02/11/17) and no internal reject found during production for the same issue. Trending Period: Sep 25, 2016 to Oct 10, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and twelve other complaints reported for the product family/different lid style. Root Cause: The case quantity for 8989PG2 is 5 eaches/case. The packaging operation is manual. A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement. | 0702097123 | 0 | Y | United States | Closed | According to the reporter, on September 25, 2017, during use, the case of the device had missing two lids. There was no patient involvement. | St. Claire Medical Center | 0004307914 | 2017-09-25 | 2017-09-28 | 0 | 2017-09-25 | During use | 2017-11-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989PG2, Serial/Lot No:175963, Description:8989PG2 18 GAL CHEMO YEL PG2 HINGE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID Model Number: 8989PG2 Serial or Lot Number: 175963 ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The case of sharps containers was missing 2 lids Below are the details for Product - Model No:V8980, Serial/Lot No:V8980, Description:V8980 Sharps Container IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for V8980 Sharps Container Model Number: V8980 Serial or Lot Number: V8980 ------------------------------------------------------------------------ ----------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The case of sharps containers was missing 2 lids PHYSICIAN LAST NAME: Trent PHYSICIAN FIRST NAME: Kathy PHYSICIAN CLINIC: St. Claire Medical Center PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Morehead PHYSICIAN STATE CODE: KY PHYSICIAN POSTAL CODE: 40351 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura ADDITIONAL COMMENTS ABOUT THE EVENT: In each case of these sharps containers, 2 lids were missing. | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-25 | 0702097123-20 | V8980 | V8980 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | VHA PLUS SHRPS 8GL RED LG VL10 | MITG - Sharps Disposal | 002927 | V8980 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product V8980 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E34363 (08/06/17 to 08/08/17) and no internal reject found during production for the same issue. The lot number was obtained from the complaint attachment photo. Trending Period: Sep 25, 2016 to Oct 10, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family/different lid style. Root Cause: The case quantity for V8980 is 10 eaches/case. The packaging operation is manual. A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement. | 0702097123 | 0 | Y | United States | Closed | According to the reporter, on September 25, 2017, during use, the case of the device had missing two lids. There was no patient involvement. | St. Claire Medical Center | 0004307914 | 2017-09-25 | 2017-09-28 | 0 | 2017-09-25 | During use | 2017-11-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989PG2, Serial/Lot No:175963, Description:8989PG2 18 GAL CHEMO YEL PG2 HINGE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8989PG2 18 GAL CHEMO YEL PG2 HINGE LID Model Number: 8989PG2 Serial or Lot Number: 175963 ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The case of sharps containers was missing 2 lids Below are the details for Product - Model No:V8980, Serial/Lot No:V8980, Description:V8980 Sharps Container IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for V8980 Sharps Container Model Number: V8980 Serial or Lot Number: V8980 ------------------------------------------------------------------------ ----------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The case of sharps containers was missing 2 lids PHYSICIAN LAST NAME: Trent PHYSICIAN FIRST NAME: Kathy PHYSICIAN CLINIC: St. Claire Medical Center PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Morehead PHYSICIAN STATE CODE: KY PHYSICIAN POSTAL CODE: 40351 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura ADDITIONAL COMMENTS ABOUT THE EVENT: In each case of these sharps containers, 2 lids were missing. | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-26 | 0702098788-10 | 8901SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8901SA 1PT RED CONTAINER BLOOD NEEDLE | MITG - Sharps Disposal | 002927 | 8901SA |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8931PG2 evaluated for the missing lids/missing containers. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number provided is unknown. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 25, 2016 to Oct 11 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8901SAis 100 eaches/case. The complaint is only for 1 each. There is no other information in the complaint regarding if the customer received the full case or only one unit. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as missing lids product should not be used. | 0702098788 | 0 | Y | United States | Closed | The customer states that the lids are missing from delivery. | CLOCK MEDICAL SUPPLY | 0013144558 | 2017-09-25 | 2017-09-26 | 0 | 2017-09-25 | Unknown | 2017-11-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Rose Matthews | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-26 | 0702099949-10 | 8906 | 00177059 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 8906 YEL 1 QRT CONT PHLEBOTOMY X100 | MITG - Sharps Disposal | 002927 | 8906 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8906 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 177059 (05/25/17 to 06/02/17) and no internal reject found during production. Trending Period: Sep 26, 2016 to Oct 17, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 8906 is 100 eaches/case. The lid box is placed in the finished good case first before containers per standard work. Also there is a weight scale for check. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702099949 | 0 | Y | Canada | Closed | According to the reporter, on September 26, 2017, during use, they received a case without lids. There was no patient involvement. | C.H.U.Q ARMAND-VIAU | 0012516257 | 2017-09-26 | 2017-09-28 | 0 | 2017-09-26 | During use | 2017-11-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Daigle, Billy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Nous avons commandé le produit # 8906 sur ma commande 2340311 par contre , nous avons reçu une caisse de 100 contenants mais il n'y avait pas de couvercle à l'intérieur . Lot #(10) 00177059 . Below are the details for Product - Model No:8906, Serial/Lot No:00177059, Description:8906 YEL 1 QRT CONT PHLEBOTOMY X100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8906 YEL 1 QRT CONT PHLEBOTOMY X100 Model Number: 8906 Serial or Lot Number: 00177059 ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event They received a case of 8906 without lids SALESREP USER ID: daiglb1 SALESREP EMP ID: 322696 SALESREP NAME: Daigle, Billy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daigle, Billy ADDITIONAL COMMENTS ABOUT THE EVENT: Field contact Catherine Hallé Acheteur / secteur inventaire Approvisionnements et gestion contractuelle Direction de la logistique CHU de Québec - Université Laval Centre administratif 775, rue Saint-Viateur, bureau 129 Québec (Québec) G2L 2Z3 Téléphone: 418-525-4444 poste 86186 Télécopieur: 418-692-1754 catherine.halle@chuq.qc.ca As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-26 | 0702100045-10 | 8970Y | ASKU | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | CRACKED CONTAINER | AFC | KC:CRACKED CONTAINER | The Product 8970Y evaluated for the cracked container. The reported condition cannot be observed as no photo or no physical sample a vailable. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 20, 2016 to Oct 16, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product. Root Cause: The lot number is unknown so DHR cannot be reviewed. As lot number was unknown it cannot be determined when the product was manufactured. The packaging improvement was implemented as per CDP-0005979 in June, 2016 per CAPA2015-010078 for 8970Y. As there is no lot number available it cannot be determined whether the product is made prior to the June 2016. The potential root cause may be container have been cracked during shipping and handling due to external stress/forces manufactured prior to June 2016 packaging improvement change. Also it was noted in the complaint there was no safety or harm to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. The packaging changes were implemented as per CAPA2015-010078 for 8970Y. | 0702100045 | 0 | Y | Canada | Closed | Customer reports that during vaccine clinic noticed the entire bottom of the sharps container cracked open. Exposed to blood and potential needle stick injury. | COCHRANE HEALTH CENTER | 0013157339 | 2017-09-20 | 2017-09-27 | 0 | 2017-09-07 | During use | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Emily Recupero | 0 | 0 | 0 | 0 | Complete | Yes
2017-09-27 | 0702100085-10 | 89651 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 89651 2GAL RED CONT H DROP CLR LIDX20 | MITG - Sharps Disposal | 002927 | 89651 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING WELD | The Product 89651 evaluated for the missing weld. The reported condition was observed during sample evaluation. The cowl and lid sample was missing weld operation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 26, 2016 to Nov 10, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to follow lid assembly welding operation step during manufacturing before packaging lids in to the case. The lot number is unknown so DHR review cannot be performed. The Standard work for assembly and inspection procedure were reviewed and found adequate for manufacturing process. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event as unassembled lid should not be used. There is no known impact or consequence to patient or end user safety. | 0702100085 | 0 | Y | Japan | Closed | The customer reports the lid came off form the container prior to use. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-09-26 | 2017-09-28 | 0 | 2017-09-21 | Prior to Use | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Janice Nevius | Gatha Kothari | 43049 | 43049 | The sample received is only lid WP8965N. The lid and cowl component were received without welding operation. As from sample evaluation it was observed that lid assembly step was missed during manufacturing operation. The cowl and lid missed the weld operation for the assembly. | Complete | Yes
2017-09-27 | 0702103705-10 | 8998PG2 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998PG2 18 GAL RED PG2 HINGE LID X5 | MITG - Sharps Disposal | 002927 | 8998PG2 |  |  |  | RFR | PACKAGING | AFC | KC:PACKAGE WRONG COMPONENTS | The Product 8998PG2 evaluated for the wrong component-lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot number provided is unknown. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 27, 2016 to Oct 18, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the lid type prior to seal the case. The case quantity for 8998PG2 is 5 eaches/case. The complaint quantity is 7 eaches which shows the case was broken down outside of Cardinal Health. Also there is no lot number information provided so further analysis cannot be performed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as wrong component product should not be used. | 0702103705 | 0 | Y | United States | Closed | According to the reporter, during pre operation, on September 27, 2017, sliding lids instead of hinged lids were sent to customer. There was no patient involvement. | MCLAREN PORT HURON | 0012542911 | 2017-09-27 | 2017-09-28 | 0 | 2017-09-27 | Prior to Use | 2017-11-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8998PG2, Description:8998PG2 18 GAL RED PG2 HINGE LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8998PG2 18 GAL RED PG2 HINGE LID X5 Model Number: 8998PG2 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Sliding Lids instead of Hinged Lids were sent to customer Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Kivel PHYSICIAN FIRST NAME: Denise PHYSICIAN CLINIC: McLaren Port Huron PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 1221 PINE GROVE AVENUE PHYSICIAN CITY: Port Huron PHYSICIAN STATE CODE: MI For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-09-29 | 0702108046-10 | 8982 | 17E33063 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8982 2GAL YEL CONTAINER CHEMO X20 | MITG - Chemo Protection Safety | 003209 | 8982 |  |  |  | RFR | LOW PUNCTURE FORCE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 8982 evaluated for the contaminated needle stick. The reported condition was not observed as no actual incident product physical sample available for product analysis. The samples sent were unused container/lids which met acceptance criteria. DHR Review was performed and met the specification. DHR Review: Device History Records data reviewed for 17E33063 (07/31/17 to 08/03/17) and no internal reject found during production. Trending Period: Sep 29, 2016 to Oct 26, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the same failure for the product family. Root Cause: The potential root cause may be customer misuse. There is no information provided for where the needle was punctured through the container. Also there is no information provided what part of the body nurse got needle stick. As sample or photo was not available so it is unknown what location container was punctured? The IFU states that “Use additional caution with container that do not have handles.” The most probable root cause is failure to follow the instruction for use per IFU for carrying/lifting the sharps disposal. The actual physical sample with the needle stick was not available for product analysis. The sample received were unused containers/lids which met specification criteria per DHR review. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The received samples were inadequate to perform further product analysis. The actual incident sample was not sent by customer. As per complaint there was no known impact or consequence to patient. | 0702108046 | 0 | Y | United States | Closed | The customer reports needle was poking through the container and it stabbed a nurse. | Cancer Institute of Doylestown Hospital | 0016566348 | 2017-09-29 | 2017-09-29 | 0 | 2017-09-27 | During use | 2018-02-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Non-Serious Illness Injury | 0 | Emily Recupero | Gatha Kothari | 43075 | 43075 | The sample parts received in the Cardinal overshipper case. There are total 11 containers and 12 lids received. The product was not used. The complaint is for the “Needle Stick”. The actual sample was not received so product analysis cannot be performed. The reported condition cannot be observed for product analysis. | Complete | Yes
2017-09-29 | 0702108133-10 | 8507SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | LID DOES NOT ATTACH / STAY ATTACHED | AFC | KC:LID FIT ISSUE | The Product 8507SA evaluated for the lid fit issue. The reported condition cannot be observed as no photo or no physical sample available for evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 29, 2016 to Oct 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The root cause cannot be determined as no sample or photos available for product analysis. Also DHR review cannot be performed as lot no. is unknown. The potential cause may be user error to follow guideline for lid fit per IFU (Instruction for Use). Also as per the complaint quantity 6 eaches shows the finished good case 20 eaches/case was broken down outside of Cardinal Health. Which may cause not using correct lid and container together. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. Based on the above information and the minimal information present in the complaint and lack of sample/photo defect was not observed. | 0702108133 | 0 | Y | United States | Closed | The customer reports lid were not fitting. Customer said the tops of the Sharps container have a snap on them and the container does not have snaps | Henry Schein | 0014601769 | 2017-09-29 | 2017-09-29 | 0 | 0 | Unknown | 2017-12-13 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Emily Recupero | 0 | 0 | 0 | 0 | Complete | Yes
2017-09-29 | 0702108885-10 | 8881907102 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881907102 10ML ORAL MEDICATION SYR | MITG - Needles & Syringes | 003208 | 8881907102 |  |  |  | RFR | SHEATH/CAP ISSUE | AFC | KC:MISSING COMPONENTS | Note Creator: Jeremy Roberts Note Date: 11/03/17 Note Description: Addition of Investigation Summary. Incident Description: According to the reporter, on September 29, 2017, during pre operation, the box of the devices had no caps. There was no patient involvement. Product Analysis: As of 11/03/17 no samples or pictures have been provided. Without samples, we are unable to confirm the reported condition. If samples are received, the complaint may be re-opened and the investigation continued. DHR Review: A review of the device history record was unable to be performed without a lot number. Typical lot size for part 8881907102 is 288 cases (144,000 pieces) and production line 7 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Root Cause Analysis: The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. The assessment eval uated potential causes: Man, Material, Method, Measurement, Machine, and Mother Nature (Environment). The significant causes are noted and addressed in the corrective and preventative action section as appropriate. Man Causes: Operator Training (Not Significant – Operators are trained to performed required functions in accordance with packaging instructions provided in the design specification). Measurement Causes: Visual Acceptance Criteria (Not Significant – Criteria exists and is required to be used during manufacturing. The AQL assigned with this defect is 1.5. In the event random sampling is conducted under normal LII sampling plans the customer would acc on 2/rej on 3. AQL has been met for the reported condition of 1 case missing tips). Material Causes: - Tips (Not Significant – Material is confirmed during lot start up activities to ensure product is being packaged as expected.) Method Manual Packaging Process (Significant – The customer confirmed one case of missing tip caps). The AQL assigned to this failure mode is 1.5. Acceptance criteria has been met. Further processing at customer location (Significant – A bulk package of tips is placed in the over shipper carton with the bulk package of oral syringes. It is possible the tips fell out while handling at the customer location. ) Machine Causes: N/A: The related packaging process is completed manually. Environment Causes: N/A: Production is completed in a controlled manufacturing environment. Root Cause: In the event of a failure during the manufacturing process, the most probable root cause is human error. It is not clear from the information provided if the tips may have fallen out/been misplaced while handling at the customer’s site. Trend Analysis: GS10000125 is used for trending analysis in support of GS10005190 to help determine when a formal investigation may be necessary. A review of customer complaints as well as internal NCR rejects pertaining to this failure mode (Missing Components) and product family was performed. Zero other incidents of this nature for this product family was received in the trending period from 9.29.16 – 11.3.17. Zero internal NCR rejects of this nature for this product family were received in the trending period from 6.29.17 – 11.03.17. Formal Investigation, Corrective and Preventive Action: Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Risk Management Assessment: RE00060107 Rev. 2. pFMEA for Syringe Focus Factory Manual Packaging was reviewed which confirmed the failure and root cause for “missing component” is reflected and accounted for. | 0702108885 | 0 | Y | United States | Closed | According to the reporter, on September 29, 2017, during pre operation, the box of the devices had no caps. There was no patient involvement. | VITALITY MEDICAL | 0012894249 | 2017-09-29 | 2017-10-02 | 0 | 2017-09-29 | Prior to Use | 2017-11-18 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Gill, Chris PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Sanket Patel of the facility ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881907102, Description:8881907102 10ML ORAL MEDICATION SYR IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Vitality Medical Inc. CONTACT NAME AND PHONE: Ashok Shah 801-733-4449 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881907102 10ML ORAL MEDICATION SYR Model Number: 8881907102 Serial or Lot Number: null ------------------------------------------------------------------------ ----------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Asked but unknown Was the IFU for this product received and followed by the user? No When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Boxes of Syringes contained no caps Was the customer notified the product should be returned? Yes FACILITY NAME: Vitality Medical Inc FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 940 Clopper Road FACILITY CITY: Gaithersburg FACILITY STATE CODE: MD FACILITY POSTAL CODE: 84121 FACILITY DAY PHONE: 801-733-4449 SALESREP USER ID: gillc5 SALESREP EMP ID: 341090 SALESREP NAME: Gill, Chris PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gill, Chris | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-03 | 0702115035-10 | 8881112599 | 17F27163X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | PLUNGER ISSUE | AFC | KC:UNIDENTIFIED EMBEDDED MATTER | Note Creator: Jeremy Roberts Note Date: 11/30/17 Note Description: Addition of Investigation Summary. Reference attached/above investigation for a full concise summary of all investigations performed. A review of the device history record was completed. The review confirmed no abnormal process conditions were present during the ma nufacturing of this product that would lead to the reported condition and all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. With the contaminate being on the plunger rod, the molded di scoloration from foreign matter and inclusions is classified to be in a non-critical and non-fluid path area. As such the manufacturing process considers these as a process indicator. In the event of Molding Contaminations/Specks in non-fluid path critical areas, the AQL assigned would be 6.5. With random sampling conducted under normal LII sampling plans, the lot would be accepted on 18/rejected on 19. The method of root cause analysis that was implemented in this investigation was to conduct a lab analysis of the complaint sample p rovided, perform a process walk and complete a Six M assessment. The most probable root cause has been identified as personnel placing the resin throat magnet assembly on/next to the yellow ladder during cleaning. This caused paint chip contaminates to adhere to the magnet assembly and upon reinstallation contaminates making way into the resin stream pre-molding. Lab analysis of the complaint and a paint chip sample confirmed a match in primary elemental composition. Based on the information available and the investigation findings, additional actions are deemed unnecessary at this time. The mold startup process was updated and released after production of the reported lot to specify (1) number of molding shots required to verify correct operation of equipment/process and (2) number of consecutive molding shots required to verify process stabilization. This update will help ensure any contamination which has worked its way into the material stream has been processed out prior to production. Personnel will be notified of the event during shift huddles to ensure awareness. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. | 0702115035 | 0 | Y | United States | Closed | Customer reports that 8 syringes were found with yellow color defect. Some plungers very clearly discolored and some slightly stained. | AQUABILITI | 0012728353 | 2017-10-03 | 2017-10-03 | 0 | 2017-10-03 | During use | 2017-12-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Mitchell Eichas - QA Engineer. ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Product at Aquabiliti CONTACT NAME AND PHONE: Mitchell Eichas, Ph# 615-833-2633 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event LOT# 17F27163X - 8 syringes found with yellow color defect. Some plungers very clearly discolored and some slightly stained. Pictures to be forwarded to Product Monitoring and samples have been requested. Was the customer notified the product should be returned? No Please explain Only 8 sysringes SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires a written investigation close-out letter. | Veronica Dasari | Christopher Gildea | 43032 | 43032 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2017-10-03 | 0702115036-10 | 8881112599 | 17D04863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | PLUNGER ISSUE | AFC | KC:UNIDENTIFIED OBJECT IN FLUID PATH | Note Creator: Jeremy Roberts Note Date: 11/15/17 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record (DHR) was completed. The DHR review showed that all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Product analysis was performed. Upon magnified visual examination of the rubber tip, nothing was observed to be embedded however the tip did include a .10mm2 piercing along the top rib. A yellow/orange discoloration was also observed. This discoloration surrounded two distinct areas where foreign matter may have adhered to the rubber tip surface? These areas measured a minimum of .42mm2 and .10mm2. Removal of saline solution by the customer likely compromised sample integrity. The manufacturing site has no history of metal contamination embedded in rubber tips. An investigation was performed. The team brainstormed potential root cause areas where metal contamination may be generated, insufficiently removed, and/or insufficiently detected from the production process. A production process walk was also performed. Through the investigation approximately fifteen areas/sources were identified to have higher potential for metal contamination. Samples from the sources were sent out for lab analysis. Lab analyses of potential sources were reviewed against the lab analysis provided by the customer. Results were inconclusive. A direct match did not exist nor could the potential sources be ruled out due an indirect comparison of potential metal contaminates vs. saline solution. The saline solution showed high levels of Sodium (99.97%) followed by lower levels of Zinc (0.009%), Potassium (0.008%), Calcium (0.004%), Phosphorus (0.002%), Nickel (0.001%), Iron (0.001%) and Other (0.005%). Not part of the manufacturing process, but relevant to the investigation into metals, Zinc and Nickel are commonly used for coating or plating iron to inhibit corrosion as well as provide aesthetically pleasing finishes. The most likely cause of oxidation (discoloration of saline) is trace amounts of iron. Root cause investigation could not confirm a manufacturing defect. Contaminates may be introduced through the customer filling and/or assembly operations. The following may be helpful in determining root cause and potential corrective actions at the customers site: (1) Obtaining saline solution control samples pre and post filling for comparison (2) Air flow evaluation and filtration in filling equipment (3) Use of stainless steel component parts in filling equipment (4) Maintenance program which controls tools used during maintenance activities (5) Specification of cleaning requirements following corrective maintenance activities (6) Use of double inline filtration during filling with filter challenges conducted regularly (7) Evaluate feasibility of implementing metal detection. Based on the information available, trending analysis and investigation findings, additional actions are deemed unnecessary at this time. Formal investigation CAPA2016-010882 was previously completed with numerous actions implemented after production of the reported lot. Action most notably implemented includes 100% inspection for ferrous, non-ferrous, and Stainless Steel metals. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0702115036 | 0 | Y | United States | Closed | Customer reports that one syringe had a piece of metal embedded into the plunger tip. This piece of metal is believed to cause discoloration of saline solution giving it a rusty appearance. | AQUABILITI | 0012728353 | 2017-10-03 | 2017-10-03 | 0 | 2017-10-03 | During use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Mitchell Eichas, QA Engineer - meichas@aquabiliti.com ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: No processes at Aquabiliti where metal objects are introduced to the device. Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event LOT# 17D04863X - one syringe had a piece of metal embedded into the plunger tip. This piece of metal believed to cause discoloration of saline solution giving it a rusty appearance. SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires written close-out investigation letter. Pictures will be forwarded to Product Monitoring and sample has been requested. | Veronica Dasari | Christopher Gildea | 43032 | 43032 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2017-10-04 | 0702117650-10 | 8506Y | ASKU | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8506Y YEL 5 QRT CONT S STAR C B DR X20 | MITG - Sharps Disposal | 002927 | 8506Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MALFUNCTIONING LID | The Product 8506Y evaluated for the malfunction lid. The reported condition cannot be observed as no photo or no physical sample available to evaluate door functionality. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Sep 08, 2016 to Oct 06, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family/lid type from the same customer. Root Cause: The ampule is very light weight sharps disposal which required door function “LIFT TO ASSURE DISPOSAL” as printed on the door. It may require different style of lid-container to be used in this type of setting based on type of sharps disposal (ampule is light weight). The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. It may be different style lid/container opening unit should be used for customer’s specific end use application. Also as per instruction for use needs to be follow for door function “LIFT TO ASSURE DISPOSAL”. | 0702117650 | 0 | Y | Canada | Closed | The customer states that ampules are getting stuck in a sharps container lid. The customer had to flip the lid 2- 3 times to get the ampule pieces to drop in to the container. | Surrey Memorial Hospital | 0001097356 | 2017-09-08 | 2017-10-04 | 0 | 0 | During use | 2017-12-08 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Janice Nevius | 0 | 0 | 0 | 0 | Complete | Yes
2017-10-06 | 0702120776-10 | 1522SA | 17F02663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING OR DAMAGED LABEL | The Product 1522SA evaluated for the missing or damaged label. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17F02663 (07/12/17 to 07/20/17) and no internal reject found during production. Trending Period: Oct 5, 2016 to Oct 30, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure verify the label prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing/damaged label product should not be used. | 0702120776 | 0 | Y | Japan | Closed | The customer reported no labeling were on the products. The bottom half of the labeling were cut off. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-10-05 | 2017-10-06 | 0 | 2017-10-04 | Prior to Use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-06 | 0702120903-10 | 8938 | 17F14863 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | CONTAINER ISSUE | AFC | KC:COSMETIC ISSUE | The Product 8938 evaluated for the cosmetic issue. The reported condition was not observed as no physical sample or photo available for product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17F14863 (7/28/17 to 08/03/17) and no internal reject found during production. Trending Period: Oct 6, 2016 to Nov 3, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The scratch is Minor A class per GS10000913. As per inspection standard it has only reaction plant for correction unless scratch depth is measureable per specification. There is no photo or physical sample provided for the product analysis. So further analysis cannot be performed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk as product meet current GS10000913 criteria for scratch per DHR review, therefore does not pose any harm to the end user for product use. | 0702120903 | 0 | Y | Japan | Closed | The customer reported that the product had a scratch on the bottom. No patient involvement. | Niti-on Co., Ltd | 0015877231 | 2017-10-06 | 2017-10-06 | 0 | 2017-10-05 | Prior to Use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-06 | 0702121558-10 | 1522SA | 176903 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 1522SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176903 (05/24/2017 to 05/31/2017) and no internal reject found during production. Trending Period: July 13, 2016 to Oct 12 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the product. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 1522SA is 60 eaches/case. The complaint is only for 1 each. As per complaint quantity it seems like the case was broken outside of Cardinal Health for distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. There was no patient involvement. | 0702121558 | 0 | Y | United States | Closed | According to the reporter, on July 13, 2017, there were missing lids on the devices. There was no patient involvement. | HEMA SOURCE INC | 0013080249 | 2017-07-13 | 2017-10-09 | 0 | 2017-07-13 | Unknown | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Park, Lynn Event Context: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Our customer Hemasource ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Missing lids. Customer requests replacement. Below are the details for Product - Model No:1522SA, Serial/Lot No:176903, Description:1522SA 2 QRT RED CONTAINER PHLEBOTOMY IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 1522SA 2 QRT RED CONTAINER PHLEBOTOMY Model Number: 1522SA Serial or Lot Number: 176903 ------------------------------------------------------------------------ ---------------------------------- Please provide a detailed description of the event Missing lids. Customer requests replacement SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Park, Lynn As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-06 | 0702122697-10 | 8881112059 | 17E02863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | CONTAMINATION | AFC | KC:PLUNGER BROKE | Note Creator: Jeremy Roberts Note Date: 12/1/17 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record (DHR) was completed. The DHR showed that there were some challenges during manufacturing, with cavity 49A being closed. It is important to note that all acceptance criteria checks per established sampling levels were within predefined limits. Product analysis was completed based on pictures supplied. The reported condition was observed. Cavity specific molding defects are a good indication of mold damage. The method of root cause analysis that was implemented in this investigation was to conduct a Five Why assessment that evaluated cau ses. The most probable root cause identified stemmed from a pick and place robot which is used for removing mold runners from the injection molding machine during processing. One of the operating cylinders of the robot likely intermittently ceased up which caused the robot gripper to strike cavity 49A and damage it. Based on the information available, no trend established and investigation findings, additional actions are deemed unnecessary at th is time. Work order CL-2566662 was completed to replace the cylinder after production of the reported lot. This Minor A defect is considered damage, functional less than desired (1.5 AQL). Acceptance criteria per established levels are within acceptable limits. If additional information is received warranting further analysis, the investigation may be resumed? This complaint will be used for tracking and trending purposes. | 0702122697 | 0 | Y | United States | Closed | According to the reporter, on October 6, 2017, the device has defective plunger rod 12 mL, also it has black specs on it. | SMITHS MEDICAL ASD INC | 0012875719 | 2017-10-06 | 2017-10-09 | 0 | 2017-10-06 | During use | 2017-12-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Toussaint, Douglas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Serial/Lot No:17E02863X, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No WHO HAS POSSESSION: Field Contact REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: 17E02863X ------------------------------------------------------------------------ ----------------------------------- Please provide a detailed description of the event defective plunger rods. Was the customer notified the product should be returned? Yes SALESREP USER ID: toussd1 SALESREP EMP ID: 237108 SALESREP NAME: Toussaint, Douglas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Toussaint, Douglas As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-10 | 0702126230-10 | 8938 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938 RED 18 GAL CONT W SLIDING LID X5 | MITG - Sharps Disposal | 002927 | 8938 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:DAMAGED PRODUCT | The Product 8938 evaluated for the damaged product. The reported condition cannot be observed as no photo or no physical sample available for evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Oct 10, 2016 to Nov 6, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The root cause cannot be determined as no sample or photos available for product analysis. Also DHR review cannot be performed as lot no. is unknown. The complaint event description states that “the product was distorted prior to its use.” The potential root cause may be product may have been damaged during shipping and handling due to external stress/forces which may cause lid does not attach/stay attached. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. Based on the above information and the minimal information present in the complaint and lack of sample/photo defect was not observed. | 0702126230 | 0 | Y | Japan | Closed | Customer reports that the lid was not able to close because the product was distorted prior to its use. No Patient involved. | Tsuchiura Kyodo General Hospital | 0008028674 | 2017-10-10 | 2017-10-10 | 0 | 2017-10-09 | Prior to Use | 2018-02-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-10-10 | 0702127848-10 | 31144010 | UNK | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31144010 5 QRT TORTOUS PATH TRANS RED | MITG - Sharps Disposal | 002927 | 31144010 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31144010 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Oct 06, 2016 to Nov 1, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family same customer. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 31144010 is 30 eaches/case. The box is designed to hold 30 containers at the bottom layer and 30 lids on top layer. As per the complaint, the Cardinal Health case may be broken down for further distribution to the end user by the 3rd party distribution. There is no information in the complaint regarding whether the product was received in the original case or not. This is more repacking of product for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702127848 | 0 | Y | United States | Closed | Customer reports that order should have been 30 containers and lids but came with 57 containers and 24 lids. | Holy Family Memorial Med Ctr | 0015922514 | 2017-10-06 | 2017-10-11 | 0 | 0 | Prior to Use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-11 | 0702130273-10 | 31144010 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31144010 5 QRT TORTOUS PATH TRANS RED | MITG - Sharps Disposal | 002927 | 31144010 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31144010 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Oct 05, 2016 to Oct 31, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 31144010 is 30 eaches/case. The box is designed to hold 30 containers at the bottom layer and 30 lids on top layer. As per the complaint, the Cardinal Health case may be broken down for further distribution to the end user by the 3rd party distribution. There is no information in the complaint regarding whether the product was received in the original case or not. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702130273 | 0 | Y | United States | Closed | The Customer reported: 57 containers were received and 24 lids. They were expecting 30 containers with 30 lids. | Holy Family Memorial Inc | 0015739014 | 2017-10-05 | 2017-10-11 | 0 | 2017-10-05 | Prior to Use | 2017-11-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-20 | 0702149778-10 | 8989 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Oct 20, 2016 to Nov 10, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and thirteen other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 8989 is 5 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702149778 | 0 | Y | United States | Closed | According to the reporter, during use, the device were missing from the case. There was no patient involvement. | FRANCISCAN HEALTH INDIANAPOLIS | 0013213370 | 2017-10-20 | 2017-10-23 | 0 | 0 | During use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Lids were missing from the case of SC - Item # 8989 Below are the details for Product - Model No:8989, Description:8989 18 GAL CONT CHEMO HNGD LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8989 18 GAL CONT CHEMO HNGD LID X5 Model Number: 8989 Serial or Lot Number: null ------------------------------------------------------------------------ --------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids were missing from the case of SC - Item # 8989 Was the customer notified the product should be returned? Yes For Physician with Siebel ID 0007220555: PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-20 | 0702150173-10 | 8991 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8991 RED 18GAL SHARPS DISP X5 | MITG - Sharps Disposal | 002927 | 8991 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8991 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Oct 20, 2016 to Nov 10, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and thirteen other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 8991 is 5 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702150173 | 0 | Y | United States | Closed | According to the reporter, on October 17, 2017, during use, the device were missing from the case of product. There was no patient involvement. | ST JOSEPH MERCY HOSPITAL CHELSEA | 0012552045 | 2017-10-20 | 2017-10-24 | 0 | 2017-10-20 | During use | 2017-11-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Ceneri, Jessica PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8991, Description:8991 RED 18GAL SHARPS DISP X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8991 RED 18GAL SHARPS DISP X5 Model Number: 8991 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lids Missing for the case of product Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Newberg PHYSICIAN FIRST NAME: Art PHYSICIAN COUNTRY CODE: USA PHYSICIAN CITY: Chelsea PHYSICIAN STATE CODE: MI For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: cenerj1 SALESREP EMP ID: 279536 SALESREP NAME: Ceneri, Jessica PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Ceneri, Jessica As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-10-23 | 0702153667-10 | 8881512852 | 17B00363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:COMPONENT BROKE IN USE | Note Creator: Jeremy Roberts Note Date: 12/06/17 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was completed. The review confirmed no abnormal process conditions were present during the ma nufacturing of this product that would lead to the reported condition and all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Product analysis was unable to be completed. As of 12/5/17 no samples or pictures have been provided. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated causes. Root cause investigation could not confirm a manufacturing defect as this may occur if excess angular stress is placed on the tip during use or a defect/malfunctioning mating device damaged the tip causing premature failure/breakage. Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. The reported condition could not confirm a manufacturing defect. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0702153667 | 0 | Y | Canada | Closed | According to the reporter, on August 16, 2017, during use, the device's tip snapped off/sprayed. No harm resulted to nurse. There was no patient involvement. | St. Paul's Hospital - SK | 0001097465 | 2017-10-23 | 2017-10-24 | 0 | 2017-08-16 | During use | 2018-02-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | MITG: UNKNOWN | 0 | FIRST NOTIFIED NAME: Gutheil, Amy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512852, Serial/Lot No:17B00363X, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: 17B00363X ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Report States:"Nurse inserted slip tip syringe into blood transfer device - RN reports syringe tip snapped off/sprayed. No harm resulted to nurse - PPE was worn. Another product issue report has been sent in for the blood transfer device. I have a sample of the product and will send it to Mike." Was the customer notified the product should be returned? Yes For Physician with Siebel ID 0016401720: PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Attn: Ronalda Henderson, Manager of Clinical Projects Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 SALESREP USER ID: guthea1 SALESREP EMP ID: 252384 SALESREP NAME: Gutheil, Amy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gutheil, Amy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0016401720 Alara Procyshen cannot find a BP in GCH | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-10-26 | 0702160534-10 | 8933 | 17E09463 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8933 RED 12 GAL CONT W/HINGED LID X10 | MITG - Sharps Disposal | 002927 | 8933 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8933 evaluated for the missing lids. The reported condition cannot be observed as the sample received is only one container in non-Cardinal Health Overshipper product case. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E10963 (06/01/17 to 06/03/17) and no internal reject found during production. Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case. The case quantity for 8933 is 10 eaches/case. As per complaint attachment “FW CASE-2017-00019880 TYCO HEALTHCARE K” 4 eaches are missing lids for Lot number 17E10963. There is no information in the complaint regarding whether the customer received the product in original Cardinal Health Overshipper case or not. As per the attached report the product was from distributor. The case was broken outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702160534 | 0 | Y | United States | Closed | The customer states that we are down to only 1 container that has a lid, the other 9 containers are missing lids | CAMPBELL SURGERY CENTER | 0013172082 | 2017-10-26 | 2017-10-26 | 0 | 2017-10-25 | Unknown | 2017-11-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | Gatha Kothari | 43053 | 43053 | One container is received without lid in the non-Cardinal Health Overshipper product case. The reported condition stated by the customer could not be determined as product was broken down into smaller quantity by distributor. | Complete | #N/A
2017-10-26 | 0702160534-20 | 8933 | 17E10963 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8933 RED 12 GAL CONT W/HINGED LID X10 | MITG - Sharps Disposal | 002927 | 8933 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8933 evaluated for the missing lids. The reported condition cannot be observed as the sample received is only one container in non-Cardinal Health Overshipper product case. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E10963 (06/01/17 to 06/03/17) and no internal reject found during production. The lot number was verified with the complaint sample received. Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case. The case quantity for 8933 is 10 eaches/case. As per complaint attachment “FW CASE-2017-00019880 TYCO HEALTHCARE K” 1 each- missing lid for Lot number 17E10963. There is no information in the complaint regarding whether the customer received the product in original Cardinal Health Overshipper case or not. As per the attached report the product was from distributor. The case was broken outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702160534 | 0 | Y | United States | Closed | The customer states that we are down to only 1 container that has a lid, the other 9 containers are missing lids | CAMPBELL SURGERY CENTER | 0013172082 | 2017-10-26 | 2017-10-26 | 0 | 2017-10-25 | Unknown | 2017-11-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | Gatha Kothari | 43053 | 43053 | One container is received without lid in the non-Cardinal Health Overshipper product case. The reported condition stated by the customer could not be determined as product was broken down into smaller quantity by distributor. | Complete | #N/A
2017-10-26 | 0702160534-30 | 8933 | 00175628 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8933 RED 12 GAL CONT W/HINGED LID X10 | MITG - Sharps Disposal | 002927 | 8933 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8933 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 175628 (03/08/17 to 03/09/17) and no internal reject found during production. Trending Period: Oct 26, 2016 to Nov 14, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case. The case quantity for 8933 is 10 eaches/case. As per complaint attachment “FW CASE-2017-00019880 TYCO HEALTHCARE K” 4 eaches are missing lids for Lot number 175628. There is no information in the complaint regarding whether the customer received the product in original Cardinal Health Overshipper case or not. As per the attached report the product was from distributor. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702160534 | 0 | Y | United States | Closed | The customer states that we are down to only 1 container that has a lid, the other 9 containers are missing lids | CAMPBELL SURGERY CENTER | 0013172082 | 2017-10-26 | 2017-10-26 | 0 | 2017-10-25 | Unknown | 2017-11-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | Gatha Kothari | 43053 | 43053 | One container is received without lid in the non-Cardinal Health Overshipper product case. The reported condition stated by the customer could not be determined as product was broken down into smaller quantity by distributor. | Complete | #N/A
2017-10-30 | 0702165370-10 | 89651 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 89651 2GAL RED CONT H DROP CLR LIDX20 | MITG - Sharps Disposal | 002927 | 89651 |  |  |  | RFR | LID ISSUE | AFC | KC:INCORRECT ASSEMBLY | The Product 89651 evaluated for the incorrect assembly. The reported condition was observed in the photo evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E11063 (06/15/17 to 06/18/17) and no internal reject found during production. Trending Period: Oct 30, 2016 to Nov 21, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to follow lid assembly step during manufacturing before packaging lids in to the case. This was resulted in the incorrect orientation of lid assembly. There was no specification change to the product. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The incorrect mis-assembled lid should not be used. There is no known impact or consequence to patient or end user safety. | 0702165370 | 0 | Y | Japan | Closed | The customer stated that the lid on the product is seemed to change, it had no plate parts. They asked if there had been specification change. | Niti-on Co., Ltd | 0015877231 | 2017-10-30 | 2017-10-30 | 0 | 2017-10-27 | Prior to Use | 2017-11-28 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | Gatha Kothari | 43060 | 43060 | The photo received is only lid WP8965N. In the photo 89651 was shown from two different lot number 175852 and 17E11063. The part from lot number 175852 was used to compare two lid assembly. The WP8965N from lot number 175852 is as per specification. The lid assembly WP8965N from lot number 17E11063 seems like incorrect orientation of cowl to lid during assembly. | Complete | #N/A
2017-11-01 | 0702171528-10 | 8970Y | 00173849 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | CONTAINER ISSUE | AFC | KC:COSMETIC ISSUE | The Product 8970Y evaluated for the cosmetic issue(bowed bottom). The reported condition was observed in the photo/physical sample for product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 173849 (11/07/16 to 11/10/17) and no internal reject found during production. Trending Period: Oct 30, 2016 to Nov 17, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The DHR was reviewed and there were no issues found during production. The probable root cause may be water temperature of the mold during production or process. The human error and failure to identify bowed bottom during visual inspection; which is based on AQL. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with minimal risk bowed bottom container must be discarded and should be replaced for patient use. There was no patient involvement per complaint. | 0702171528 | 0 | Y | Canada | Closed | Customer reports that bottom of container is bulging making it unstable. | OKOTOKS COMMUNITY HEALTH CENTRE | 0012536922 | 2017-10-30 | 2017-11-07 | 0 | 2017-10-03 | Unknown | 2018-01-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | Gatha Kothari | 43056 | 43056 | No physical sample was submitted with this complaint. The reported condition stated by the customer was observed in “photo 2”. The c ontainer bottom was bowed. The container may rock or unstable when placed on a flat horizontal surface. | Complete | #N/A
2017-11-03 | 0702177177-10 | 8881512852 | 16KD4863X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO IDENTIFIABLE MALFUNCTION | AFC | KC:DIFFICULT TO OPEN PACKAGE | Note Creator: Jeremy Roberts Note Date: 12/13/17 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was unable to be completed as the lot number provided was invalid. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Product analysis was unable to be completed. As of 12/13/17 no samples or pictures have been provided. A defect could not be determined at this time. Torque testing is performed throughout the lot to ensure opening specifications are met. Customer complaint onfirmed attempting only method 1A (twisting the cap) from the supplied instructions for use. It is advised that if the cap does not remove from twisting (method 1A), that pressing the edge of the cap above the heat stake with the thumb or on a tapping on a hard surface be used to break the seal (methods 1B and/or 1C). When the seal is broken, the cap should twist or pull off with applying minimum pressure. It should be noted that removal of the cap will be difficult if squeezing the edges of the cap while attempting to remove it. Based on the information available and the investigation findings, a formal investigation is not deemed necessary at this time. The reported condition could not confirm a manufacturing defect. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0702177177 | 0 | Y | Canada | Closed | The customer reported that the syringes case will not open to release syringe. Customer also reported that the there was a "Code 2: Minor injury to patient or staff. Sore hands & wrists" | ROYAL UNIVERSITY HOSPITAL | 0012598710 | 2017-11-02 | 2017-11-13 | 0 | 2017-07-01 | During use | 2017-12-21 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Gutheil, Amy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Reported as "Code 2: Minor injury to patient or staff. Sore hands & wrists" "Several personnel have tried to open these syringes by twisting and can not get them to release." 21 syringes with this lot number were affected. Below are the details for Product - Model No:8881512852, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? Yes Provide details No harm to patient, but minor harm to staff Please provide a detailed description of the event Report States: " Syringes case will not open to release syringe" For Physician with Siebel ID 0016444698: PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Please send to: Ronalda Henderson Manager of Clinical Projects Saskatoon Health Region Supply Chain Management 103 Hospital Dr. Saskatoon, SK., S7N 0W8 Ronalda.Henderson@saskatoonhealthregion.ca SALESREP USER ID: guthea1 SALESREP EMP ID: 252384 SALESREP NAME: Gutheil, Amy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gutheil, Amy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0016444698 Cathy Coote cannot find a BP in GCH | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-11-03 | 0702177190-10 | 8900SA | 8900SA | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Nov 03, 2016 to Nov 22, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and five other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8900SA is 100 eaches/case. The lid box is placed first in the overshipper case and the containers are placed after in the finished good case. Also there is a weight scale for check in manufacturing as control. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702177190 | 0 | Y | United States | Closed | the customer reported all 100 lids were missing from the case; product was not broken down beforehand | CLINCH VALLEY MEDICAL CENTER | 0012644348 | 2017-11-03 | 2017-11-07 | 0 | 2017-11-03 | Prior to Use | 2017-12-02 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Gill, Chris PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8900SA, Serial/Lot No:8900SA, Description:8900SA 1QT RED CONTAINER PHLEBOTOMYX100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 Model Number: 8900SA Serial or Lot Number: 8900SA ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event All 100 lids were missing from the case; product was not broken down beforehand PHYSICIAN LAST NAME: stillwell PHYSICIAN FIRST NAME: Danny PHYSICIAN CLINIC: Clinch Valley Medical Center PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 6801 GOV G C PEERY HWY PHYSICIAN CITY: Richlands PHYSICIAN STATE CODE: VA PHYSICIAN POSTAL CODE: 24641 PHYSICIAN DAY PHONE: 2765966702 PHYSICIAN EMAIL: albert.stillwell@lpnt.net For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: gillc5 SALESREP EMP ID: 341090 SALESREP NAME: Gill, Chris PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gill, Chris ADDITIONAL COMMENTS ABOUT THE EVENT: Lucy from Thermo Fisher Email:mailto:lucy.bojczuk@thermofisher.com As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-11-06 | 0702180278-10 | 1510SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | AUTO DROP 50 | MITG - Collection Systems | 002901 | 1510SA |  |  |  | RFR | CRACKED | AFC | KC:NO FAILURE | The Product 1510SA evaluated for the no failure. The reported condition stated by the customer could not be determined as samples or photos were not available for evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Nov 6, 2016 to Dec 04, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: The product event description is not very clear to perform investigation. The 1510SA has three components Needle holder, disposable thread and Dispenser. The Needle holder has white spring buttons assembled with it for the spring function. As per the Instruction for Use, hold Needle Holder and push straight down onto disposable thread. Once thread is engaged (snap fit no cross threads on the thread); pull holder straight up from dispenser. Applying moderate pressure, twist blood needle clockwise onto holder. To dispose of needle: Using one hand, push holder straight down into the funnel of a KendallTM sharps container to release needle and disposable thread. Note: If disposal unit with funnel is not available: Using one hand, hold unit over a sharps container opening with the needle tip pointed down. Release needle by depressing both white spring buttons (located on the sides of barrel) at the same time. (see attached photo from IFU) The needle holder is not designed to be cross threaded. The disposable thread must be used with the assembly. The samples were received have needle holder, disposable thread and dispenser. The samples received are with needles attached to it from BD part number 368608. All the samples with needle holder, thread and needles are tested. The thread is designed as per specification without thread. The thread is snap fit into the needle holder. The needle used with it has thread on it to attach to the needle holder. One of the BD needle Part 368608 has damaged thread which caused the needle to detach from the 1510SA which is Cardinal Health product. The root cause was on Non-Cardinal Health part BD needle 368608 (see photo attached). Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product event description is not very clear to perform investigation and also the lot number is unknown. The needle holder is not designed to be cross threaded. The disposable thread must be used with the assembly. As per complaint there was no consequences or impact to patient. | 0702180278 | 0 | Y | Canada | Closed | According to reporter, the needle or the tube holder is cross threaded and that is what would make it detach one from the other | 0 | 0 | 2017-11-06 | 2017-11-14 | 0 | 2017-11-06 | Intra Operative | 2018-01-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Lafortune, Dino PATIENT MEDICAL HISTORY: Asked but unknown PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Collection Systems ---------------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? No Did the event lead to or extend patient hospitalization? No Was there any injury as a result of the event? No injury If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT OUTCOME: Alive - No injury Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: Yes Below are the details for Product - Model No:1510SA, Description:AUTO DROP 50 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for AUTO DROP 50 Model Number: 1510SA Serial or Lot Number: null ------------------------------------------------------------------------ ------- When was the issue noticed? During procedure Was this a Bone Marrow product? No Please provide a detailed description of the event In a blood collection procedure on various patients, when using the device, it needs to be screwed onto a tube holder. As it was screwed on to the needle holder, the nurse would then proceed to puncture the patient and when they would insert a blood collection tube inside the tube holder, the holder would seperate from the needle. Nurse was saying that it seems like the needle or the tube holder is cross threaded and that is what would make it detach one from the other. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Lane PHYSICIAN FIRST NAME: Jennifer PHYSICIAN CLINIC: University Hospital of Northern British Columbia PHYSICIAN COUNTRY CODE: CAN PHYSICIAN ADDRESS1: 1475 Edmonton St PHYSICIAN CITY: Prince George PHYSICIAN STATE CODE: BC PHYSICIAN POSTAL CODE: V2M 1S2 PHYSICIAN DAY PHONE: 250-565-2282 For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes FACILITY NAME: University Hospital of Northern British Columbia FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 1475 edmonton street FACILITY CITY: Prince George FACILITY STATE CODE: BC FACILITY POSTAL CODE: V2M 1S2 FACILITY DAY PHONE: 250-565-2282 SALESREP USER ID: laford1 SALESREP EMP ID: 252160 SALESREP NAME: Lafortune, Dino PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Lafortune, Dino As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZFACILIT 6981031 University Hospital of Northern British Columbia cannot find a BP in GCH | Veronica Dasari | Gatha Kothari | 43117 | 43118 | The 3 samples were received have needle holder, disposable thread, dispenser and BD Needles. The samples received are with needles attached to it from BD part number 368608. All the samples with needle holder, thread and needles are tested. The thread is designed as per specification without thread. The thread is snap fit into the needle holder. The needle used with it has thread on it to attach to the needle holder. One of the BD needle Part 368608 has damaged thread which caused the needle to detach from the 1510SA which is Cardinal Health product. The defect was on Non-Cardinal Health part (BD needle 368608). See attached photo of damage thread. (Photo- BD-Damage Thread 2 and 3) | Complete | #N/A
2017-11-08 | 0702185247-10 | 31323333 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 2 GALLON RIGID PLASTIC CONTAINER ( TRANS | MITG - Sharps Disposal | 002927 | 31323333 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MALFUNCTIONING LID | The Product 31323333 evaluated for the malfunction lid. The reported condition cannot be observed as no physical sample available to evaluate door functionality. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Nov 08, 2016 to Dec 01, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: It is unknown what type of sharps were being disposed into 31323333. If sharps are very light weight; then sharps disposal which required door function “LIFT TO ASSURE DISPOSAL” as printed on the door. Also there is no information provided in the complaint regarding “how full was the container”. If the container is full above fill line than it needs to be disposed. It may require different style of lid-container for the user application based on type of sharps disposals. The root cause is determined to be customer misuse. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with no user harm or injury. It may be different style lid/container opening unit should be used for customer’s specific end use application. Also as per instruction for use; follow the instruction for door function “LIFT TO ASSURE DISPOSAL”. | 0702185247 | 0 | Y | United States | Closed | The customer states that the lid would not easily drop the sharp while being disposed. The user had to manipulate the lid to drop the sharps. | BELL DRUG | 0013177002 | 2017-11-07 | 2017-11-08 | 0 | 2017-11-07 | Prior to Use | 2018-02-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-11-08 | 0702185707-10 | 8991 | 00176815 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8991 RED 18GAL SHARPS DISP X5 | MITG - Sharps Disposal | 002927 | 8991 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8991 evaluated for the missing lids. The reported condition was observed in the physical sample provided. The case may have been broken in to small quantity at the distribution. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176815 (05/12/17 to 05/16/17) and no internal reject found during production. Trending Period: Nov 8, 2016 to Nov 28, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eleven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 8991 is 5 eaches/case. As per complaint report there was external distributor. The probable root cause might be the original case was broken down for 3 eaches and repackaged without lids to send to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702185707 | 0 | Y | United States | Closed | The customer states no lids. patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-02 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | Gatha Kothari | 43067 | 43067 | The physical sample received. The original overshipper Cardinal Health case had only 3 containers in the box. So the case was missing 2 containers and 5 lids. | Complete | #N/A
2017-11-08 | 0702185708-10 | 31143699 | 17E20763 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1 GALLON RIGID PLASTIC CONTAINER (RED) | MITG - Sharps Disposal | 002927 | 31143699 |  |  |  | RFR | LID ISSUE | AFC | KC:MOLDING ISSUE | The Product 31143699 evaluated for the molding issue. The reported condition “hard to slide” is subjective per user. The door was fu nctioning as its intended use. The warp was observed on the door may cause hard to slide. DHR Reviewed and met acceptance criteria. DHR Review: A Device History Record reviewed for Lot Number 17E20763 manufactured from 07/02/17 to 07/12/17. There was no internal reject found during production. Trending Period: Nov 7, 2016 to Dec 7, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The most probable root cause may be door was slightly warped. The warp may be due to fluctuation in water temperature for lid/door part during molding process. As per GS10000913 warp is allowed as long as part is functional for use. Also “hard to slide” is subjective per user. The door was functioning as its intended use. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event with no user harm or injury. The product met acceptance criteria for form, fit and function for its intended use. | 0702185708 | 0 | Y | United States | Closed | The customer states the lids are hard to get on and off.The patient information details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2018-01-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Gatha Kothari | 43073 | 43076 | The original overshipper case with 32 containers and lids was received. The lid fit test was performed on the product. There was no functional issue found with the lid fit. The sliding door was little hard to slide. But the sliding door was function as its intended use. The reported condition for the sliding door hard to slide was observed but this is not a failure of product as this is subjective to user. | Complete | #N/A
2017-11-08 | 0702185712-10 | 8970- | 17E32263 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970 2GAL RED CONTAINER COLLECTIONX20 | MITG - Sharps Disposal | 002927 | 8970- |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8970 evaluated for the missing lids. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E32263 (06/24/17 to 07/11/17) and no internal reject found during production. Trending Period: Nov 7, 2016 to Dec 04, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the product family/different lids. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8970 is 20 eaches/case. As per complaint report there was external distributor. As per complaint product line item the quantity is 4 eaches. The physical samples received for product analysis was in the non-Cardinal Health Overshipper case. As per the attached report and product analysis it was determined the product was from distributor. The probable root cause is case was broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702185712 | 0 | Y | United States | Closed | The customer states,the product has no lids. Patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | Gatha Kothari | 43073 | 43073 | The samples were received in the non-Cardinal Health Overshipper box. The total quantity received in the box was 4 containers. The 4 containers were for lot number 17E32263. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box which shows the product was broken down and repackaged by distributor from original Cardinal Health case. | Complete | #N/A
2017-11-08 | 0702186329-10 | 8507SA | 176916 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176916 (05/11/17 to 05/18/17) and no internal reject found during production. Trending Period: Nov 08, 2016 to Nov 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint product line item the quantity is 14 eaches. The physical samples received for product analysis was in the non-Cardinal Health Overshipper case. As per the attached report and sample analysis it was determined the product was from distributor. The case may be broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702186329 | 0 | Y | United States | Closed | Customer states the product has no lids. Patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Gatha Kothari | 43066 | 43066 | The samples were received in the non-Cardinal Health Overshipper box. The total quantity received in the box was 27 containers. The 14 containers were for lot number 176916. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box which shows the product was broken down and repackaged by distributor from original Cardinal Health case. | Complete | #N/A
2017-11-08 | 0702186329-20 | 8507SA | 17G28763 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17G28763 (09/12/17 to 09/19/17) and no internal reject found during production. Trending Period: Nov 08, 2016 to Nov 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint product line item the quantity is 4 eaches. The physical samples received for product analysis was in the non-Cardinal Health Overshipper case. As per the attached report and sample analysis it was determined the product was from distributor. The case may be broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702186329 | 0 | Y | United States | Closed | Customer states the product has no lids. Patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Gatha Kothari | 43066 | 43066 | The samples were received in the non-Cardinal Health Overshipper box. The total quantity received in the box was 27 containers. The 14 containers were for lot number 176916. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box which shows the product was broken down and repackaged by distributor from original Cardinal Health case. | Complete | #N/A
2017-11-08 | 0702186329-30 | 8507SA | 176177 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 176177 (03/11/17 to 03/18/17) and no internal reject found during production. Trending Period: Nov 08, 2016 to Nov 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint product line item the quantity is 5 eaches. The physical samples received for product analysis was in the non-Cardinal Health Overshipper case. As per the attached report and sample analysis it was determined the product was from distributor. The case may be broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702186329 | 0 | Y | United States | Closed | Customer states the product has no lids. Patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Gatha Kothari | 43066 | 43066 | The samples were received in the non-Cardinal Health Overshipper box. The total quantity received in the box was 27 containers. The 14 containers were for lot number 176916. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box which shows the product was broken down and repackaged by distributor from original Cardinal Health case. | Complete | #N/A
2017-11-08 | 0702186329-40 | 8507SA | 17E04663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17E04663 (06/02/17 to 06/09/17) and no internal reject found during production. Trending Period: Nov 08, 2016 to Nov 27, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint product line item the quantity is 4 eaches. As per complaint product line item the quantity is 4 eaches. The physical samples received for product analysis was in the non-Cardinal Health Overshipper case. As per the attached report and sample analysis it was determined the product was from distributor. The case may be broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702186329 | 0 | Y | United States | Closed | Customer states the product has no lids. Patient involvement details not provided. | HENRY SCHEIN INC DENTAL | 0012964014 | 2017-11-07 | 2017-11-08 | 0 | 0 | Unknown | 2017-12-03 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | Gatha Kothari | 43066 | 43066 | The samples were received in the non-Cardinal Health Overshipper box. The total quantity received in the box was 27 containers. The 14 containers were for lot number 176916. The reported condition stated by the customer could not be determined as samples received were in non-Cardinal Health Finished good box which shows the product was broken down and repackaged by distributor from original Cardinal Health case. | Complete | #N/A
2017-11-10 | 0702191571-10 | 8906 | 00177059 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 8906 YEL 1 QRT CONT PHLEBOTOMY X100 | MITG - Sharps Disposal | 002927 | 8906 |  |  |  | RFR | NEEDLE STICK/PUNCTURE | AFC | KC:CONTAMINATED NEEDLE STICK | The Product 8906 evaluated for the contaminated needle stick. The reported condition was observed in the sample/photo as needle was protruding in the bottom of the container. DHR Review was performed and met the specification. DHR Review: Device History Records data reviewed for 177059 (05/25/17 to 06/02/17) and no internal reject found during production. Trending Period: Nov 07, 2016 to Dec 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the same failure for the product. Root Cause: The potential root cause associated with this event is improper handling/carrying of a filled sharps container for disposal below re commended fill line by user. As per complaint the container was punctured in the bottom side. The container was full above the fill line. As per the IFU carrying/lifting instructions “Never hold or carry used container below recommended fill line. Use additional caution with containers that do not have handles.” Also per IFU precaution “do not hold or carry container close to the body”. Also there is a caution on the label stating “Forcing could cause puncture”. As from the sample the sharps were filled way above the fill line. (Photo 8906_4). As from the sample analysis it was observed that needle was forced in to the container with excessive force as the needle was bent. (Photo 8906_1). Due to excessive force it was resulted in the container puncture. From the sample analysis all the sharps were disposed into the container needle facing downwards. It seems like customer may need different style Sharps Container for their specific application in which sharps can be disposed horizontally. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The container was filled way above the fill line. The excessive force may be used to drop sharps in the container resulted in the container puncture. Also the carrying/lifting instructions were not followed by user. As from product analysis it seems like customer may need different style Sharps Disposal product for their specific application in which it allows to drop sharps horizontally rather than vertically. | 0702191571 | 0 | Y | Canada | Closed | The customer reports that a nurse suffered a needle stick injury because it was protruding from the sharps container. | CENTRE DE SANTE MEMPHREMAGOG | 0013070124 | 2017-11-10 | 2017-11-21 | 0 | 2017-11-07 | During use | 2018-02-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | Serious Illness Injury | FIRST NOTIFIED NAME: Daigle, Billy PATIENT MEDICAL HISTORY: Unknown, but will follow up for more information PATIENT INJURY DETAILS: The nurse will be follow with a post exposition protocol. This nurse was stung by the needle. PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- Was medical or surgical intervention needed to prevent a permanent impairment of a function? Unknown Unknown, but will follow-up for more information Did the event lead to or extend patient hospitalization? Unknown Unknown, but will follow-up for more information Was there any injury as a result of the event? Unknown Unknown, but will follow-up for more information If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : Unknown, but will follow up for more information PATIENT AGE: Unknown, but will follow up for more information PATIENT WEIGHT: Unknown, but will follow up for more information PATIENT OUTCOME: Alive - With injury Complaint Source: Field Contact ASSOCIATED WITH PATIENT: Yes PRODUCT EVENT DETAILS OTHER: This nurse is follow by a post exposition protocol to biological liquids by a needle from unknow source. Below are the details for Product - Model No:8906, Serial/Lot No:00177059, Description:8906 YEL 1 QRT CONT PHLEBOTOMY X100 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 8906 YEL 1 QRT CONT PHLEBOTOMY X100 Model Number: 8906 Serial or Lot Number: 00177059 ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? Yes Provide details Nurse Please provide a detailed description of the event A 25FR. needle pierced the sharp container on the side. When the nurse bring this sharp container in the dirty utility she was stung by a 25 FR. needle. The sharp container wasn't full and was closed easily. Was the customer notified the product should be returned? Yes SALESREP NAME: Hoyt, Julie PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daigle, Billy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | Gatha Kothari | 43077 | 43077 | The sample photos were available for product analysis. The needle was protruding through container in the bottom part of the container. From the photo it seems like needle was forced in to the container with excessive force as the needle was bent in the photo. Due to excessive force it was resulted in the container puncture. | Complete | Yes
2017-11-21 | 0702210025-10 | 8881112083 | 17B04363X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112083 SYR 112 S-PLK SP X800 | MITG - Needles & Syringes | 003208 | 8881112083 |  |  |  | RFR | PLUNGER ISSUE | AFC | KC:PLUNGER BENT | Note Creator: Jeremy Roberts Note Date: 1/11/18 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was completed. The review confirmed no abnormal process conditions were present during the ma nufacturing of this product that would lead to the reported condition and all acceptance criteria inspections per established sampling levels were within acceptable limits during the production process. Product analysis was completed based on the sample supplied. The plunger rod was observed to have numerous bends (damage functional less than desired) on the flights near the rubber tip. This is considered a minor defect with a 1.5 AQL. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated causes. Root cause identified the most probable cause to be Inherent Variation in the aidlin material transfer process during manufacturing (Based on the quantity reported being defective as 1 and trending results). Based on the information available, no trend established and investigation findings, additional actions are deemed unnecessary at th is time. Acceptance criteria per established levels are within acceptable limits. If additional information is received warranting further analysis, the investigation may be resumed? This complaint will be used for tracking and trending purposes. | 0702210025 | 0 | Y | United States | Closed | The customer reported: Syringe shows signs of being melted. | DEROYAL IND INC | 0013005301 | 2017-11-20 | 2017-11-21 | 0 | 0 | Unknown | 2018-02-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | Jeremy Roberts | 43110 | 43110 | Note Creator: Jeremy Roberts Note Date: 1/10/18 Note Description: Addition of analysis. Product Analysis: One syringe was returned for evaluation. The syringe was not in originally supplied packaging. Upon visual inspection of the syringe subassembly, the plunger rod was observed to have damage (functional less than desired) on the plunger flights near the rubber tip. The plunger rod sample was confirmed to be manufactured on mold 6455, cavity 35A. | Complete | #N/A
2017-11-22 | 0702213400-10 | 31143665 | 1045534864 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31143665 10 GAL RED CONTAINER SHARPSX6 | MITG - Sharps Disposal | 002927 | 31143665 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31143665 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Nov 21, 2016 to Dec 08, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 31143665 is 6 eaches/case. As per complaint description it seems operator error to verify required amount of lid prior to seal the case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702213400 | 0 | Y | United States | Closed | The customer reported: There are 3 EA out of the 6 EA that she ordered for ITEM PN31143665 have no covers. | PATIENT PARTNERS SURGERY CTR | 0013228248 | 2017-11-21 | 2017-11-22 | 0 | 2017-11-21 | Unknown | 2017-12-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2017-11-26 | 0702216430-10 | 31144010 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31144010 5 QRT TORTOUS PATH TRANS RED | MITG - Sharps Disposal | 002927 | 31144010 |  |  |  | RFR | LID ISSUE | AFC | KC:MALFUNCTIONING LID | The Product 31144010 evaluated for the malfunction lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Nov 17, 2016 to Dec 8, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: As lot number is unknown DHR cannot be reviewed. The packaging change was implemented as per CAPA2014-008259 which became effective as of 06/23/2015. Refer to CAPA2014-008259 for complete root cause analysis. As lot number is unknown the potential root cause may be the product was manufactured prior to the packaging change. Formal Investigation: Based on complaints trend analysis CAPA2014-008259 was opened for further investigation. Refer to CAPA for more details for correction and containment. The packaging change implemented and CAPA2014-008259 is closed. Conclusion: This issue has been determined to be an isolated event as it was discovered prior to use. The locked lid should not be used and must be replaced. | 0702216430 | 0 | Y | United States | Closed | Customer reports that all of the tops were already closed and locked and unable to be used. | 0 | 0 | 2017-11-17 | 2017-11-26 | 0 | 2017-11-17 | Prior to Use | 2017-12-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-11-27 | 0702217352-10 | 31323333 |  | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 2 GALLON RIGID PLASTIC CONTAINER ( TRANS | MITG - Sharps Disposal | 002927 | 31323333 |  |  |  | RFR | LID ISSUE | AFC | KC:MALFUNCTIONING LID | The Product 31323333 evaluated for the failure to drop sharp-malfunction lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number was provided. Trending Period: Nov 21, 2016 to Dec 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The probable root cause may be lid accidently locked as final lock during use by user. If the sharps container is already full above the fill line than also door would not open. Also there is no photo or sample available for product analysis. The information in the complaint is minimal to perform investigation for the failure without sample. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. As per complaint there was no patient involvement. | 0702217352 | 0 | Y | United States | Closed | Customer reported that the the lid would not open. | BELL DRUG | 0013177002 | 2017-11-21 | 2017-11-27 | 0 | 0 | Prior to Use | 2017-12-20 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-11-29 | 0702223660-10 | 8939CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 18 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8939CRS |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:LID FIT ISSUE | The Product 8939CRS evaluated for the lid fit issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Nov 28, 2016 to Jan 12, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The most probable root cause per CAPA2017-012361 is the R7413B container tool has a larger rim height than the R7413A tool, which causes the lid locks to be difficult to secure around the rim. As per lid tool CAPA 2016-010397: 1) A product drawing DR10001522 was created to build the 1st tool (R7417A). At some point in time, the core in this tool was adjusted to improve the functional requirement of lid fit. But the product drawing was never updated to reflect those adjustments. The same, outdated drawing was used to build the 2nd tool (R7417B). 2) During DESIGN VERIFICATION RE00023792, the lid fit & function was only tested on containers produced from one container mold, when the lid is expected to perform with containers produced from several other molds in the fleet. When the functional requirement of lid fit was tested with containers produced from these other molds, it was found that product from the 2nd tool (R7417B) failed to comply with the functional requirement of lid fit. Also as per IFU “Snap lid on container by pressing down on all sides. Listen for clicks. Please check that the container has suffered no shipping damage before placing into use.” So customer needs to check lid was snapped on all sides before further use. As lot number was not provided unable to determine the manufacturing date for the product code to know if it is manufactured prior to listed CAPA implementation. Formal Investigation: Refer to CAPA2017-012361 and CAPA2016-010397 for more details for corrective and preventive action for R7413B container tool. Both CAPA are already implemented corrective action as planned. As for this complaint the lot number is unknown it cannot be determined when was the product manufactured. Conclusion: Refer to CAPA2017-012361 and CAPA2016-010397 for more details for corrective and preventive action. | 0702223660 | 0 | Y | Canada | Closed | The customer reported that the lid removed from the container. | CJD CHUQ Hotel Dieu de Quebec | 0001221545 | 2017-11-28 | 2017-12-07 | 0 | 0 | Intra Operative | 2018-01-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Daigle, Billy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8939CRS, Description:18 GAL CHEMO RED SLIDE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 18 GAL CHEMO RED SLIDE LID Model Number: 8939CRS Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event When the transporter took this full container to bring it in his truck the lid removed from the container. No injury in this case, but they had to used scotch tape to close the container. Was the customer notified the product should be returned? Yes SALESREP NAME: Roy, Jocelyn PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daigle, Billy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-11-29 | 0702223879-10 | 8938Y | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8938Y YEL 18 GAL SHARP CONT SLIDINGLID | MITG - Sharps Disposal | 002927 | 8938Y |  |  |  | RFR | LID ISSUE | AFC | KC:NO FAILURE | #N/A | 0702223879 | 0 | N | Canada | Closed | The customer reported that they doesn't like the lid of the product. There is confusion when it is on a pedal cart. The workers think this is a chemo container because of the red lid. | *HOSPITAL LAVAL | 0012883876 | 2017-11-29 | 2017-12-07 | 0 | 2017-11-29 | Intra Operative | 2017-12-14 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Daigle, Billy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8938Y, Description:8938Y YEL 18 GAL SHARP CONT SLIDINGLID IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8938Y YEL 18 GAL SHARP CONT SLIDINGLID Model Number: 8938Y Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The customer doesn't like the lid of this product. There is confusion when it is on a pedal cart. The workers think this is a chemo container because of the red lid. Was the customer notified the product should be returned? Yes SALESREP USER ID: daiglb1 SALESREP EMP ID: 322696 SALESREP NAME: Daigle, Billy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daigle, Billy ADDITIONAL COMMENTS ABOUT THE EVENT: Mélanie Paquette hygiène et salubrité As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-04 | 0702231326-10 | 8991 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8991 RED 18GAL SHARPS DISP X5 | MITG - Sharps Disposal | 002927 | 8991 |  |  |  | RFR | CART ISSUE | AFC | KC:DAMAGED PRODUCT | An investigation of the reported condition began 01/12/2018 by Tollin McCane. The investigation is entered in GCH by Gatha Kothari on 01/15/2018. The Product 8991 evaluated for the damage product. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: December 04, 2016 – January 15, 2018 Trending Results: A review of internal inspection rejects for broken/damaged components was performed as well as a review of customer complaints covering the trending period. Results showed 1 other complaint reported for broken/damaged components in this product family. Internal inspection rejects were evaluated and no trend was found for this failure type this product family. Root Cause: Product code 8991 is a Red, 18 Gallon Sharps Container. This container is designated for use in the 8941FP or 8992H carts. Product code 8991 is not shipped in combination with either compatible cart. The potential root cause has been identified as damage during shipping. As per complaint RFR code is “Cart Issue”. The complaint is initiated for the 8991 which is not cart part number as stated above. These issues have been determined to be isolated events with minimal risk as damaged lids and/or containers should not be used. Correction and Containment: Based on the existing controls, examination and review of similar production samples, the internal reject and the complaint history review, additional correction or containment activities are not warranted at this time. Corrective and Preventive Action: Based on the information available and the investigation findings, a Corrective and Preventive Action (CAPA) is not deemed necessary at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Conclusion: there is no patient involvement as per customer complaint data. These issue has been determined to be isolated events with minimal risk as damaged lids and/or containers should not be used. | 0702231326 | 0 | Y | United States | Closed | The customer reported: When customer received new cart it came with a broken piece on the actual cart so customer can’t use. | Great Lakes Surgical Suites, LLC | 0016589424 | 2017-12-04 | 2017-12-04 | 0 | 2017-11-23 | Unknown | 2018-01-17 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-04 | 0702231394-10 | 8617RC | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 18 GAL RCRA HINGE TOP | MITG - Sharps Disposal | 002927 | 8617RC |  |  |  | RFR | NO CONSEQUENCES OR IMPACT TO PATIENT | AFC | KC:LID FIT ISSUE | The Product 8617RC evaluated for the lid fit issue. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Dec 04, 2016 to Jan 15, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family/lid type. Root Cause: The 8617RC 18 G RCRA Container with Hinge Lid. The most probable root cause is failure to follow guideline for lid fit per Instruction for use by customer. As per IFU “Snap lid on container with hinged lid open. Listen for snaps. Please check that the container has suffered no shipping damage before placing into use.” Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. As per complaint there was no patient involvement. | 0702231394 | 0 | Y | United States | Closed | The customer states the lids on the sharps containers are not securing. | UNIV OF KENTUCKY HOSPITAL | 0013106278 | 2017-12-04 | 2017-12-08 | 0 | 2017-12-04 | During use | 2018-02-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8617RC, Description:18 GAL RCRA HINGE TOP IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 18 GAL RCRA HINGE TOP Model Number: 8617RC Serial or Lot Number: null ------------------------------------------------------------------------ ---------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event SC lids are not securing and are partially popping off PHYSICIAN LAST NAME: Riggs PHYSICIAN FIRST NAME: Debbie PHYSICIAN CLINIC: University of Kentucky Hospital PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 1000 S. Limestone PHYSICIAN CITY: Lexington PHYSICIAN STATE CODE: KY PHYSICIAN POSTAL CODE: 40536 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: University of Kentucky Hospital FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 1000 S. Limestone FACILITY CITY: Lexington FACILITY STATE CODE: KY FACILITY POSTAL CODE: 40536 FACILITY DAY PHONE: 859-323-1276 SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-12-05 | 0702234050-10 | 8618RC | 17E30163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8618RC 18 GA RCRA HAZARDOUS WASTE X5 | MITG - Sharps Disposal | 002927 | 8618RC |  |  |  | RFR | PACKAGING | AFC | KC:MISSING ABSORBENT PADS | The Product 8618RC evaluated for the missing pads/labels (kit). The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: A Device History Record reviewed for the lot number 17E30163 (07/29/17 to 07/31/17). There was no internal reject found during production. Trending Period: Dec 5, 2016 to Dec 11, 2017 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: A probable root cause may be human error and failure to verify the amount of the absorbent pads packed prior to sealing the case. The absorbent pads (5 eaches) are kitted with labels as one kit per case. The 8618RC has 5 eaches per case. As per the complaint the quantity is 1 eaches which shows the product case was broken down outside of Cardinal Health may resulted in missing absorbent pads/labels. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk for missing absorbent pads. There is no patient involvement, so there is no risk to patient safety or harm. | 0702234050 | 0 | Y | United States | Closed | The customer reported: The container did not come with the sticker and absorbent pads. | Kona Community Hospital | 0001154821 | 2017-12-05 | 2017-12-05 | 0 | 0 | Unknown | 2017-12-30 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-05 | 0702234756-10 | 31144010 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31144010 5 QRT TORTOUS PATH TRANS RED | MITG - Sharps Disposal | 002927 | 31144010 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | Existing/Previous Investigation | Existing/Previous Investigation | 0702234756 | 0 | Y | United States | Closed | Customer reports that all of the tops were already closed and locked and unable to be used. | MEDLINE INDUSTRIES, INC | 0013144141 | 2017-11-17 | 2017-12-05 | 0 | 0 | Unknown | 2018-01-24 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-06 | 0702237411-10 | 8970- | 17J16163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970 2GAL RED CONTAINER COLLECTIONX20 | MITG - Sharps Disposal | 002927 | 8970- |  |  |  | RFR | NO INFORMATION | AFC | KC:MISSING LIDS | The Product 8970 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. The product was repackaged by distributor. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17J16163 (10/01/17 to 10/05/17) and no internal reject found during production. Trending Period: Dec 11, 2016 to Jan 25, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8970 is 20 eaches/case. As per complaint product line item the quantity is 1 each. The probable root cause is case was broken and repackaged outside of Cardinal Health for further distribution. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702237411 | 0 | Y | United States | Closed | The customer reported: missing qty 2 lids. | 0 | 0 | 2017-12-06 | 2017-12-06 | 0 | 0 | Unknown | 2018-02-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-09 | 0702243562-10 | 8934CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8934CRS |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:LID FIT ISSUE | The Product 8934CRS evaluated for the lid fit issue. The reported condition cannot be observed on physical sample during product analysis. The lid fit was performed with slight difficulty which is not a failure. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 156904 (04/14/14 to 04/17/14) and no internal reject found during production. Trending Period: Dec 7, 2016 to Jan 18, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product family. Root Cause: The most probable root cause per CAPA2017-012361 is the R7413B container tool has a larger rim height than the R7413A tool, which causes the lid locks to be difficult to secure around the rim. Formal Investigation: Refer to CAPA2017-012361 for more details for corrective and preventive action for R7413B container tool. The CAPA is already implemented corrective action as planned. The product was manufactured in year 2014. The CAPA was implemented with correction in year 2017. Conclusion: Refer to CAPA2017-012361 for more details for corrective and preventive action. | 0702243562 | 0 | Y | Canada | Closed | According to reporter, the lid is very difficult to snap on completely and the lid won't be properly secured and creates potential risk if container was ever to tip over. | St Pauls Hospital | 0008037463 | 2017-12-07 | 2017-12-19 | 0 | 2017-12-07 | During use | 2018-02-16 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Dahlbeck, Michelle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8934CRS, Description:12 GAL CHEMO RED SLIDE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 12 GAL CHEMO RED SLIDE LID Model Number: 8934CRS Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Lid is very difficult to snap on completely and clinicians are worried that the entire lid won't be properly secured and creates potential risk if the container was ever to tip over. I did suggest carts to help secure but currently these containers sit on the ground beside stations. Was the customer notified the product should be returned? Yes PHYSICIAN LAST NAME: Carr PHYSICIAN FIRST NAME: Angela PHYSICIAN CLINIC: St. Paul's Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Vancouver PHYSICIAN STATE CODE: BC For Physician : PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : acarr@providencehealth.bc.ca SALESREP USER ID: dahlbm1 SALESREP EMP ID: 331345 SALESREP NAME: Dahlbeck, Michelle PRIMARY CONTACT: Someone else PRIMARY CONTACT OTHER: Dahlbeck, Michelle (Prepared By) COMPLAINT OWNER NAME: Dahlbeck, Michelle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | Gatha Kothari | 43118 | 43118 | The physical sample was available for product analysis. The sample has R7413B container with R7417A lid. The lid fit was slightly difficult but all lid lugs were attached on the container. There is no risk of lid coming off. | Complete | #N/A
2017-12-11 | 0702246028-10 | 8970- | 17H13563 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970 2GAL RED CONTAINER COLLECTIONX20 | MITG - Sharps Disposal | 002927 | 8970- |  |  |  | RFR | NO INFORMATION | AFC | KC:PACKAGE WRONG COMPONENTS | An investigation of the reported condition began 01/19/2018 by Tollin McCane. The investigation is entered in GCH by Gatha Kothari on 01/23/2018. Product code 8970 has been evaluated for Package Wrong Components. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. A DHR Review has been performed with no nonconforming events identified. DHR Review: Device History Records data has been for product code 8970 lot 17H13563. This review has shown that production was conducted 09/20/17 to 09/29/17 with no nonconforming events identified. All inspections passed. Trending Period: December 11, 2016 – January 23, 2018 Trending Results: A review of internal inspection rejects for nonfunctional lid was performed as well as a review of customer complaints covering the trending period. Results showed no other complaint reported for the potential of Package Wrong Components in this product family. Internal inspection rejects were evaluated and no trend was found for this failure type this product family. Root Cause: Product code 8970 is a Red, 2 Gallon Sharps Container which employs a see-thru lid and lock with rotor attached (part number WP8979). The potential root cause has been identified as a package wrong component outside the manufacturing facility. This issue has been determined to be an isolated event with minimal risk as package wrong component product should not be used. Correction and Containment: Based on the existing controls, examination and review of similar production samples, the internal reject and the complaint history review, additional correction or containment activities are not warranted at this time. Corrective and Preventive Action: Based on the information available and the investigation findings, a Corrective and Preventive Action (CAPA) is not deemed necessary at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Conclusion: there is no patient involvement as per customer complaint data. These issue has been determined to be isolated events with minimal risk as package wrong component product should not be used. | 0702246028 | 0 | Y | United States | Closed | Customer reported: That it has wrong lid. It does not have rotor lid. | Baylor Research Institute | 0009209438 | 2017-12-11 | 2017-12-11 | 0 | 0 | Unknown | 2018-02-02 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-11 | 0702246415-10 | 31156550 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | SAG4807 SHARPS-A-GATOR 7 GAL | MITG - Sharps Disposal | 002927 | 31156550 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MALFUNCTIONING LID | An investigation of the reported condition began 01/15/2018 by Tollin McCane. The investigation is entered in GCH by Gatha Kothari on 01/16/2018. Product code 31156550 has been evaluated for malfunction lid. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: December 11, 2016 – January 16, 2018 Trending Results: A review of internal inspection rejects for nonfunctional lid was performed as well as a review of customer complaints covering the trending period. Results showed no other complaints reported for malfunction lid in this product family. Internal inspection rejects were evaluated and no trend was found for this failure type this product family. Root Cause: Product code 31156550 is a Red, 7 Gallon Sharps Container which employs a hinged lid and door assembly (part number WP4810). This lid and door assembly is designed with 2 separate locking mechanisms for use when the container is full. First, the torturous lid (small lid atop lid assembly) is equipped with 3 individual locking hooks that fit into the lid assembly, and second, the front half of the lid is locked onto the container by pushing down on each of 8 hooks along the rim. Locking of the lid remains possible even in the event of a missing hook. The potential root cause has been identified as a missing locking hook on the lid and door assembly. These issues have been determined to be isolated events with minimal risk as damaged/nonfunctional lids and/or containers should not be used. Corrective and Preventive Action: Based on the information available and the investigation findings, a Corrective and Preventive Action (CAPA) is not deemed necessary at this time. If additional information is received warranting further analysis, the investigation will be resumed. This complaint will be used for tracking and trending purposes. Conclusion: There is no patient involvement as per customer complaint data. These issue has been determined to be isolated event with minimal risk as malfunction lids should not be used. | 0702246415 | 0 | Y | United States | Closed | The customer states the container did not have a plastic piece that is used to keep the lid locked when the container is full. | SHARP REES - STEALY CORP | 0012662649 | 2017-12-11 | 2017-12-19 | 0 | 2017-11-27 | During use | 2018-01-19 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Daniel, Alan-Anthony PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Purchasing Tech ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Order was placed by Owens & Minor for the customer. Order details are below. Order # - 28218241 PO # - 81828 Item # - 31156550 (2 cases) Order date - 11/22/2017 Below are the details for Product - Model No:31156550, Description:SAG4807 SHARPS-A-GATOR 7 GAL IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for SAG4807 SHARPS-A-GATOR 7 GAL Model Number: 31156550 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Customer received their 2 cases of item #31156550 and in one case, 4 containers did not have a plastic piece that is used to keep the lid locked when the container is full. New container is not needed but replacement lids have been requested by the customer. SALESREP USER ID: daniea17 SALESREP EMP ID: 349087 SALESREP NAME: Daniel, Alan-Anthony PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daniel, Alan-Anthony As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-12 | 0702248504-10 | 8881512852 | 16MO2763X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:UNIDENTIFIED OBJECT IN FLUID PATH | Note Creator: Jeremy Roberts Note Date: 1/15/18 Note Description: Addition of investigation summary. Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record (DHR) was completed. The review reveled challenges during the manufacturing process while concluding that all acceptance criteria checks per established sampling levels were still within predefined limits prior to final release. Product analysis was unable to be completed as samples were unavailable for review. Without samples or pictures, the reported condition could not be observed. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated causes. The most probable cause was identified as inherent vibration during manufacturing leading to printer entrance dial guide rail screws/bolts loosening over time which created syringe barrel jams/scratches and ultimately the potential for particulate. Based on the information available, no trend established and investigation findings, additional actions are deemed unnecessary at th is time. Maintenance Work order CL-2454527 was completed to address the most probable root cause. Acceptance criteria per established levels are within acceptable limits. If additional information is received warranting further analysis, the investigation may be resumed? This complaint will be used for tracking and trending purposes. | 0702248504 | 0 | Y | United States | Closed | According to customer, the syringe interior coated with white powdery material. | 0 | 0 | 2017-12-12 | 2017-12-21 | 0 | 2017-12-12 | Prior to Use | 2018-02-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | No Injury | FIRST NOTIFIED NAME: Anderson, Meredith PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Veterinary clinic ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512852, Serial/Lot No:16MO2763X, Description:12cc ls IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 12cc ls Model Number: 8881512852 Serial or Lot Number: 16MO2763X ------------------------------------------------------------------------ ----------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Syringe interior coated with white powdery material. Was the customer notified the product should be returned? Yes FACILITY NAME: KALENZI PATRICK DVM FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 625 Flatiron Marketplace FACILITY CITY: Broomfield FACILITY STATE CODE: CO FACILITY POSTAL CODE: 80021 FACILITY DAY PHONE: 7206694200 SALESREP USER ID: anderm56 SALESREP EMP ID: 239999 SALESREP NAME: Anderson, Meredith PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Anderson, Meredith ADDITIONAL COMMENTS ABOUT THE EVENT: Customer Contact Name Dr. Kalenzi or Stacy Customer Phone 7206694200 As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Daniel Tamul | 0 | 0 | 0 | 0 | Complete | Yes
2017-12-13 | 0702250971-10 | 8881512878 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:EXCESSIVE LUBRICANT | Note creator: Daniel Callahan Note Date:1/19/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was unable to be performed, lot # was not provided. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Product Analysis was performed. Coating on syringes was not observed as reported. It was determined by QE at the manufacturing facility that was most likely observed by customer was silicone oil. Failure reported was replicated at facility using 12mL like barrel, silicone, and adhesive material. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). Insufficient procedure for RAM cleaning was found to be the probable cause. Corrective Action was implemented on 06/02/17 per QC023506. With an average complaint event to manufacturing production date of 5.5 months for hard-pack product codes, it can be inferred the Lot # received by customer was produced before Change Order #QC023506 was implemented. If additional information is received warranting further analysis, the investigation may be resumed. This complaint will be used for tracking and trending purposes. | 0702250971 | 0 | Y | Canada | Closed | The customer reports that the 12mL and 20mL Monoject syringes have a coating on them that prevents medication stickers from adhering. | PASQUA HOSPITAL | 0012880907 | 2017-12-11 | 2017-12-22 | 0 | 2017-11-22 | During use | 2018-02-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Gutheil, Amy PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881512878, Description:8881512878 12CC S-LK SYR ONLY O/P X480 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8881512878 12CC S-LK SYR ONLY O/P X480 Model Number: 8881512878 Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Report States: "12mL and 20mL Monoject syringes have coating on them that prevents medication stickers from adhering" Was the customer notified the product should be returned? Yes Below are the details for Product - Model No:8881520657, Description:8881520657 20ML S LK SYRINGE IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8881520657 20ML S LK SYRINGE Model Number: 8881520657 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? During use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Report States: "12mL and 20mL Monoject syringes have coating on them that prevents medication stickers from adhering" Was the customer notified the product should be returned? Yes For Physician with Siebel ID 0013764788: PHYSICIAN COMMUNICATION REQUESTED: Yes Notify Health Care Professional Address Other : Please send report to: Clinical Project Manager Wascana Rehabilitation Centre 2180 – 23rd Avenue Regina, Saskatchewan S4S 0A5 Saskatchewan Health Authority 1-306-766-5837 Linda.mccann@saskhealthauthority.com Please copy the rep when sending investigation results! SALESREP USER ID: guthea1 SALESREP EMP ID: 252384 SALESREP NAME: Gutheil, Amy PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Gutheil, Amy As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes ZINIREP 0013764788 Jo Nelson cannot find a BP in GCH | Veronica Dasari | Juan Diaz | 43109 | 43109 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2017-12-19 | 0702262051-10 | 8881512852 | 17K33063X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:EMPTY SECONDARY PACKAGE | Note creator: Daniel Callahan Note Date: 1/25/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was performed for Lot# 17K33063X, with no anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881512878 is 360 cases (480 parts per case), and production line 23 was used. No sample was returned. A picture was available as seen below/Attached in GCH showing an un-damaged sealed container, with reportedly zero product inside. The reported complaint could not be replicated at the automated packaging assembly cell on Line 23. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). Mishandling of containers during clearing of auto packer jams/picker faults due to human error was found to be the probable cause. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition presents a low risk level. This complaint will be used for tracking and trending purposes. | 0702262051 | 0 | Y | United States | Closed | The customer reported: They opened the case to pull out a box of syringes, one of the boxes is completely empty. It is still sealed | ANIMAL HEALTH INTERNATIONAL INC | 0013196324 | 2017-12-15 | 2017-12-19 | 0 | 2017-12-14 | Prior to Use | 2018-01-31 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Veronica Dasari | Daniel Callahan | 43126 | 43126 | Note creator: Daniel Callahan Note Date:1/25/18 Note Description: Addition of Analysis Product Analysis: No sample was returned. A picture was available as seen below/Attached in GCH. Based on the picture provided below, the container is still sealed and has not been opened. The integrity of the container is intact with no visual evidence of damage. The packaging of the product is 100% automated. The containers are formed, labeled, filled, and sealed without handling by operator. At the case sealing station on the production line, Glue 34-25 HM160 is applied to case and sealed. QE attempted to re-create sealing of empty case on production line with no success. Normal system controls/processes were manually overridden to index a completely empty container to the sealing station robot. The sealing station robot repeatedly damaged empty container and faulted due to lack of syringe material reinforcing the structural integrity of container. As a result, the reported complaint could not be replicated. | Complete | #N/A
2017-12-21 | 0702268140-10 | 8970- | 17J16163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970 2GAL RED CONTAINER COLLECTIONX20 | MITG - Sharps Disposal | 002927 | 8970- |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | Existing/Previous Investigation | Existing/Previous Investigation | 0702268140 | 0 | Y | United States | Closed | The customer reports the lids were not included in the case. | 0 | 0 | 2017-12-18 | 2017-12-27 | 0 | 2017-12-18 | Prior to Use | 2018-01-23 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Margolis, Cory PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Field Contact ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: The lids were not in the case and would like them to be replaced Below are the details for Product - Model No:8970-, Serial/Lot No:17J16163, Description:8970 2GAL RED CONTAINER COLLECTIONX20 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8970 2GAL RED CONTAINER COLLECTIONX20 Model Number: 8970- Serial or Lot Number: 17J16163 ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Asked but unknown When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event The lids were not included in the case SALESREP USER ID: margoc2 SALESREP EMP ID: 297918 SALESREP NAME: Margolis, Cory PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Margolis, Cory ADDITIONAL COMMENTS ABOUT THE EVENT: lids need to be sent out and replaced As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2017-12-28 | 0702276707-10 | 8970- | 17H13563 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970 2GAL RED CONTAINER COLLECTIONX20 | MITG - Sharps Disposal | 002927 | 8970- |  |  |  | RFR | PACKAGING | AFC | Existing/Previous Investigation | Existing/Previous Investigation | 0702276707 | 0 | N | United States | Closed | The customer reported: Received with wrong lids. Spoke with vendor | Baylor Research Institute | 0009209438 | 2017-12-28 | 2017-12-28 | 0 | 0 | Unknown | 2018-02-02 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-03 | 0702283188-10 | 8998S | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8998S 18GAL RED CONT SLIDE SEAL LIDX5 | MITG - Sharps Disposal | 002927 | 8998S |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MISSING LIDS | The Product 8998S evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 3, 2017 to Jan 29, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and ten other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to seal the case. The case quantity for 8998S is 5 eaches/case. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702283188 | 0 | Y | United States | Closed | The customer states upon receiving their 18 gallon sharps containers, two lids were missing in their shipment. The sample is still in use. | South Peninsula Hospital | 0015734922 | 2018-01-03 | 2018-01-06 | 0 | 2017-12-26 | During use | 2018-02-01 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Daniel, Alan-Anthony PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8998S, Description:8998S 18GAL RED CONT SLIDE SEAL LIDX5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8998S 18GAL RED CONT SLIDE SEAL LIDX5 Model Number: 8998S Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Upon receiving their 18 gallon sharps containers, two lids were missing in their shipment. PHYSICIAN LAST NAME: Soto PHYSICIAN FIRST NAME: Maria PHYSICIAN CLINIC: South Peninsula Hospital PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 4300 Bartlett St PHYSICIAN CITY: Homer PHYSICIAN STATE CODE: AK PHYSICIAN POSTAL CODE: 99603 PHYSICIAN DAY PHONE: 9072350328 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: daniea17 SALESREP EMP ID: 349087 SALESREP NAME: Daniel, Alan-Anthony PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Daniel, Alan-Anthony As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-08 | 0702290834-10 | 8970Y | 17G32663 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8970Y YEL 2 GAL CONT ROTOR LID X20 | MITG - Sharps Disposal | 002927 | 8970Y |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:CRACKED CONTAINER | The Product 8970Y evaluated for the cracked container. The reported condition was observed in the photo and sample. The container was cracked in the bottom area. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for lot number 17G32663 (09/11/17 to 09/13/17) and no internal reject found during production. Trending Period: Jan 03, 2017 to Feb 2, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and one other complaint reported for the product. Root Cause: The packaging improvement was implemented as per CDP-0005979 in June, 2016 per CAPA2015-010078 for 8970Y. The DHR was reviewed and it was packaged per specification and met acceptance criteria. The root cause cannot be determined as physical sample received was in the non-cardinal case. Also outside of Cardinal Health, other distribution center may have repackaged product. Cardinal Health has Case Quantity shipped 20 Eaches per case for 8970Y. As per the complaint no. of defective quantity is 1 each, which shows that case may have been broken down to small quantity by distribution center to ship to end user. If the product case was broken down to small quantity during distribution outside of Cardinal Health may result in external stress/forces on the container as Cardinal Health packaged was removed for further distribution. As per complaint it was empty container and the issue was found before use. So there was no safety or harm to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. The packaging changes were implemented as per CAPA2015-010078 for 8970Y. As per complaint it was empty container and the issue was found before use. So there was no safety or harm to end user. | 0702290834 |  | Y | Canada | Closed | Customer reports that bin of sharps container broken and unable to use as whole bottom cracked. Patient injury was not reported. The device will be returned for evaluation. |  |  | 2018-01-03 | 2018-01-11 |  | 2017-12-22 | During use | 2018-03-15 | N | N | N | N | N | N |  | N |  |  | Null | No Injury |  | Daniel Tamul | Gatha Kothari | 2018-03-02 | 2018-03-02 | The photos are attached in the complaint. The container was cracked at the bottom. The physical sample was received and the container was cracked at the bottom. The reported condition was observed. | Complete | Yes
2018-01-09 | 0702294291-10 | 31353611 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE | MITGSHARPS | 31353611 5 QRTG/GUARDIN-RMTRANSYEX14 | MITG - Sharps Disposal | 002927 | 31353611 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:NO FAILURE | The Product 31353611 evaluated for the no failure. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review cannot be performed as lot no. information is not available. DHR Review: Device History Records data cannot be reviewed as lot number was not provided. Trending Period: Jan 9, 2017 to Feb 2, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaint reported for the product family. Root Cause: The 31353611 is 5 Qt. GG container. The DHR cannot be reviewed as no lot number is available. As the GG lid does not have any hinge as described in the complaint description. As per attachment “reply from customer” it was stated that the customer would like to suspend this complaint due to customer would like to review their own internal process for removal and deployment of new container into the brackets. The reported condition cannot be observed or no further investigation can be performed based on limited information available. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event due to user procedure. As per complaint there was no patient involvement. | 0702294291 | 0 | Y | Canada | Closed | Customer reports that House keeping staff bruised their fingers when assembling the sharps containers and collecting them from the emergency department. The issues they reported were: 1) When they are closing the white hinged part, the 2 tabs are difficult to press to lock 2) When they go to install a new container, it is difficult to assemble the top white and bottom yellow plastic sections 3) There are problems with containers fitting properly into the holders | Surrey Memorial Hospital | 0008050653 | 2018-01-08 | 2018-01-11 | 0 | 0 | Servicing | 2018-02-07 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Prokopy, Kimberly EVENT DATE: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:31353611, Description:31353611 5 QRTG/GUARDIN-RMTRANSYEX14 IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 31353611 5 QRTG/GUARDIN-RMTRANSYEX14 Model Number: 31353611 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? Yes Provide details House keeping staff have reported to their managers of bruising to their fingers when assembling the sharps containers and collecting them from the emergency department. The issues they reported were: The issues are: 1) When they are closing the white hinged part, the 2 tabs are difficult to press to lock 2) When they go to install a new container, it is difficult to assemble the top white and bottom yellow plastic sections 3) There are problems with containers fitting properly into the holders Please provide a detailed description of the event Please see above under "Was there any harm to anyone other than the patient?" PHYSICIAN LAST NAME: Nathoo PHYSICIAN FIRST NAME: Farah PHYSICIAN CLINIC: Surrey Memorial Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Surrey PHYSICIAN STATE CODE: BC For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: prokok1 SALESREP EMP ID: 252524 SALESREP NAME: Prokopy, Kimberly PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Prokopy, Kimberly As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-10 | 0702296113-10 | 8900SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8900SA 1QT RED CONTAINER PHLEBOTOMYX100 | MITG - Sharps Disposal | 002927 | 8900SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8900SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 10, 2017 to Feb 06,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eight other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. As per the complaint reported quantity is 1 each. The case quantity for 8900SA is 100 eaches/case. This shows that the Cardinal Health case was broken into smaller quantity at the external distribution center for further distribution. This was resulted in the failure of placing proper quantity of lids by external distributor prior to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702296113 | 0 | Y | United States | Closed | The customer reported: Vendor is stating they've received item B3023-30 without lids on multiple occasions. | STEWARD HEALTHCARE | 0013184754 | 2018-01-09 | 2018-01-10 | 0 | 2017-12-27 | Unknown | 2018-02-10 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-10 | 0702296880-10 | 8991 | 00174193 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8991 RED 18GAL SHARPS DISP X5 | MITG - Sharps Disposal | 002927 | 8991 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8991 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174193 (12/06/16 to 12/12/16) and no internal reject found during production. Trending Period: Jan 9, 2017 to Feb 06, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eleven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8991 is 5 eaches/case. As per the complaint reported missing quantity is 1 each. This shows that the Cardinal Health case was broken into smaller quantity at the external distribution center for further distribution. This was resulted in the missing lids by external distributor prior to ship to end user Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702296880 | 0 | Y | United States | Closed | The customer reported they received sharp disposal containers with missing lids | Caromont Regional Medical Center | 0001109459 | 2018-01-09 | 2018-01-12 | 0 | 2018-01-09 | Prior to Use | 2018-02-09 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Hill, Lauren PATIENT EVENT DETAILS: Was a Medtronic person present at the event? Yes Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8991, Description:8991 RED 18GAL SHARPS DISP X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8991 RED 18GAL SHARPS DISP X5 Model Number: 8991 Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event "We have a case of the 8991 Red Sharps Containers that we received with no lids. The lot # is 00174193" PHYSICIAN LAST NAME: Hilburn PHYSICIAN FIRST NAME: Beth PHYSICIAN CLINIC: CaroMont PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 2511 Court Drive PHYSICIAN CITY: Gastonia PHYSICIAN STATE CODE: NC For Physician : PHYSICIAN COMMUNICATION REQUESTED: No FACILITY NAME: CaroMont FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 2511 Court Drive FACILITY CITY: Castonia FACILITY STATE CODE: NC FACILITY POSTAL CODE: 28053 SALESREP USER ID: hilll12 SALESREP EMP ID: 320674 SALESREP NAME: Hill, Lauren PRIMARY CONTACT: Someone else PRIMARY CONTACT OTHER: Hilburn, Beth (HCP) COMPLAINT OWNER NAME: Hill, Lauren As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-18 | 0702312607-10 | 8881512852 | 16F07563X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO INFORMATION | AFC | KC:COMPONENT BROKE IN USE | Note creator: Daniel Callahan Note Date: 02/15/18 Note Description: Addition of Investigation Summary A review of the device history record was performed for Lot# 16F07563X, with no anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881512852 was 406 cases (480 parts per case), and production line 23 was used on 6/29/16 – 07/01/16. Six (6) samples were returned with four (4) 18x1-1/2 A needle attachments and one (1) 16x3/4 I needle attachment. The samples received were not in their original hard-pack container and were received with evidence of further use. Syringes had been inserted into needle hubs. 2 of 4 18x1-1/2 A needles inside their sheaths were bent at a ~10 degree angle, showing evidence of applied angular force. The needles were bent at an angle >10 degrees for them to remain at ~10 degrees. All six (6) syringe luer tips were broken off and displaced inside the needle hubs. Investigator was able to replicate reported condition with a quick angular force on the attached needle (Lot#332236X, Part 8881200193). The method of root cause analysis that was implemented in this investigation was to conduct a 5 Why analysis. Excessive angular force applied on the luer tip was found to be the probable root cause. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition cannot be confirmed to be a manufacturing defect. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702312607 |  | Y | United States | Closed | The customer complaints that the tip is easily broken. |  |  | 2018-01-18 | 2018-01-18 |  |  | Unknown | 2018-03-17 | N | N | N | N | N | N |  | N |  |  | Null |  |  | Daniel Tamul | Daniel Callahan | 2018-02-07 | 2018-02-22 | Note creator: Daniel Callahan Note Date: 02/12/18 Note Description: Addition of Analysis Product Analysis: Six (6) samples were returned with four (4) 18x1-1/2 A needle attachments and one (1) 16x3/4 I needle attachment. The samples received were not in their original hard-pack container and were received with evidence of further use. Syringes had been inserted into needle hubs. Multiple samples have evidence of brownish and red fluid on the outside of the needle, showing evidence of usage on animals (per customer report section 9). Furthermore, 2 of 4 18x1-1/2 A needles inside their sheaths were bent at a ~10 degree angle, showing evidence of applied angular force. The needles were bent at an angle >10 degrees for them to remain at ~10 degrees. All six (6) syringe luer tips were broken off and displaced inside the needle hubs. Luer tips received were not cavity specific. Investigator pulled barrels (PT00052251, 88815112852) on 02/14/15 from production room, same part as damage luer tips received, and were able to replicate reported condition with a quick angular force on the attached needle (Lot#332236X, Part 8881200193). | Complete | Yes
2018-01-19 | 0702315100-10 | 8881512852 | 17K19063X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:MISSING COMPONENTS | Note creator: Daniel Callahan Note Date: 02/26/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. As of 02/26/18 no samples or pictures have been provided. Without samples or pictures, we are unable to observe the reported condition. If samples or pictures are received, the complaint may be re-opened and the investigation continued. A review of the device history record was performed for Lot# 17K19063X, with no anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881512852 was 360 cases, 28,800 parts produced with Line 23 on 10/26/17. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The most probable root cause is tech operators not trained on how to handle RAM assembly equipment during production of Spraymation test samples per LS20001213. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time. The following corrective actions were implemented: 1. Training LS20001213 added to all Tech Operators on how to prepare samples for Spraymation testing. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702315100 |  | Y | United States | Closed | The customer reported received device with plungers missing |  |  | 2018-01-17 | 2018-01-31 |  | 2018-01-17 | Unknown | 2018-03-06 | N | N | N | N | N | N |  | N |  |  | Null | No Injury | FIRST NOTIFIED NAME: Anderson, Meredith Event Context: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Central supply ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: per customer received with plungers missing Below are the details for Product - Model No:8881512852, Description:8881512852 12ML S-R SYR ONLY O/P IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Vet Clinic CONTACT NAME AND PHONE: 5305474457 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881512852 12ML S-R SYR ONLY O/P Model Number: 8881512852 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Unknown Unknown, but will follow-up for more information Was there any harm to anyone other than the patient? No Please provide a detailed description of the event per customer syringes received with no plungers Was the customer notified the product should be returned? No Please explain customer sent form by email have not connected. FACILITY NAME: Millville Veterinary Cliic FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 23653 Millville way FACILITY CITY: Millville FACILITY STATE CODE: CA FACILITY POSTAL CODE: 5305474457 SALESREP USER ID: oconnr6 SALESREP EMP ID: 238231 SALESREP NAME: O'Connor, Ryan PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Oconnor, Ryan As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-22 | 0702317443-10 | 8507SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8507SA 5 QRT CONT S/STAR C/BDR T/REDX20 | MITG - Sharps Disposal | 002927 | 8507SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8507SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 18, 2017 to Feb 12,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and six other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8507SA is 20 eaches/case. The weight scale is used as production control during manufacturing process. As per complaint product line item the quantity is 3 eaches. As per the attached report it was determined the product was from distributor. The case may be broken and repackaged outside of Cardinal Health for further distribution. This shows that the Cardinal Health case was broken into smaller quantity at the external distribution center for further distribution. This was resulted in the failure of placing proper quantity of lids by external distributor prior to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702317443 | 0 | Y | United States | Closed | Customer reports that the item came in without the lids and it cannot be used | Premier Health Associates | 0015741007 | 2018-01-18 | 2018-01-22 | 0 | 2018-01-18 | Unknown | 2018-02-15 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | 0 | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-22 | 0702317557-10 | 8537SA | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8537SA 12 QRT CONT S/STAR C/B DR T/RED | MITG - Sharps Disposal | 002927 | 8537SA |  |  |  | RFR | PACKAGING | AFC | KC:NO FAILURE | The Product 8537SA evaluated for the no failure. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 22, 2017 to Feb 20,2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the correct lids prior to ship to end user. The case quantity for 8537SA is 10 eaches/case. The event description states about photographs but there were no photos provided with the complaint. As per complaint product line item the quantity is 1 each. As per the attached report the product code is 8537SA not 8537. The case may be broken and repackaged outside of Cardinal Health for further distribution. This may have been resulted in packaging wrong lids and ship to end user. There is not enough information available for further investigation. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as there is not enough information available for further investigation. | 0702317557 |  | Y | United States | Closed | The customer reported: a case of item 8537 that had incorrect lids included. In the attached pictures you can see the lid pictured on the right is the incorrect lid that is meant to accompany the sharps container. This case should have instead included the lids pictured on the left. | PAML | 0013134547 | 2018-01-22 | 2018-01-22 |  | 2018-01-18 | Unknown | 2018-03-17 | N | N | N | N | N | N |  | N |  |  | Null |  |  | Shanika Hodge | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-25 | 0702326840-10 | 8989 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 25, 2017 to Feb 22, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and thirteen other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. As per the complaint attachment 1 case was received by customer but there is no information how many lids were missing from the case. As the information provided is very limited and there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702326840 |  | Y | United States | Closed | Customer reports that the lids were missing | Huntsville Hospital | 0011808195 | 2018-01-25 | 2018-02-06 |  | 2018-01-25 | During use | 2018-02-27 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: King, Maura PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Chemo Protection Safety --------------------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8989, Description:8989 18 GAL CONT CHEMO HNGD LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8989 18 GAL CONT CHEMO HNGD LID X5 Model Number: 8989 Serial or Lot Number: null ------------------------------------------------------------------------ --------------------------- When was the issue noticed? Before use/during setup Please provide a detailed description of the event The customer received a case of item # 8989 and the lids were missing. PHYSICIAN LAST NAME: Comer PHYSICIAN FIRST NAME: Cindy PHYSICIAN CLINIC: Huntsville Hospital PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: 101 Sivley Road PHYSICIAN CITY: Huntsville PHYSICIAN STATE CODE: AL PHYSICIAN POSTAL CODE: 35801 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: kingm23 SALESREP EMP ID: 318657 SALESREP NAME: King, Maura PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: King, Maura As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-26 | 0702329351-10 | 85021 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 85021 2QT CLEAR CONT SHARPS M/BOX X20 | MITG - Sharps Disposal | 002927 | 85021 |  |  |  | RFR | PACKAGING | AFC | KC:BROKEN CONTAINER | The Product 85021 evaluated for the broken container. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Jan 26, 2017 to Feb 23, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject no other complaint reported for the product. Root Cause: The root cause cannot be determined as no physical sample available. The potential root cause may be external stress/forces were applied on the Cardinal Health product during shipping and handling. The complaint is from distributor. As per complaint it was empty container and the issue was found before use. So there was no safety or harm to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an isolated event. The product should be inspected prior to use for any damage. As per complaint there was no safety or harm to end user. | 0702329351 |  | Y | United States | Closed | The customer reported: Sharps container plastic was broken away. The box it came in was undamaged and the item was broken in the box. | BOUND TREE MEDICAL | 0013113442 | 2018-01-26 | 2018-01-26 |  | 2018-01-22 | Unknown | 2018-03-20 | N | N | N | N | N | N |  | N |  |  | Null |  |  | Priyanka Ghule | Daniel Callahan | 43174 | 43174 | The physical sample and photograph received. One 85021 product code container and lid was received. The container was broken on one corner of the top rim. The lot number 17F11963 was obtained from sample for DHR review. | Complete | #N/A
2018-01-26 | 0702329795-10 | 8989 | 17K25163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8989 18 GAL CONT CHEMO HNGD LID X5 | MITG - Chemo Protection Safety | 003209 | 8989 |  |  |  | RFR | NO INFORMATION | AFC | KC:MISSING LIDS | The Product 8989 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17K25163 (12/04/17 to 12/10/17) and no internal reject found during production. Trending Period: Jan 3, 2017 to Feb 06, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and eleven other complaints reported for the product family. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989 is 5 eaches/case. As per the complaint reported missing quantity is 1 each. This shows that the Cardinal Health case was broken into smaller quantity at the external distribution center for further distribution. This was resulted in the missing lids by external distributor prior to ship to end user Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702329795 | 0 | Y | United States | Closed | The customer states that the sharps containers missing lid. | FROEDTERT INTEGRATED SERVICE CENTER | 0012987049 | 2018-01-03 | 2018-01-30 | 0 | 2018-01-03 | During use | 2018-02-09 | N | N | N | N | N | N | 0 | N | 0 | 0 | Null | 0 | FIRST NOTIFIED NAME: Bradley, Nicholas PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Chemo Protection Safety --------------------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: They need the lids to the containers Below are the details for Product - Model No:8989, Serial/Lot No:17K25163, Description:8989 18 GAL CONT CHEMO HNGD LID X5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, replaced by another manufacturer's product PRODUCT EVENT DETAILS: Event Details for 8989 18 GAL CONT CHEMO HNGD LID X5 Model Number: 8989 Serial or Lot Number: 17K25163 ------------------------------------------------------------------------ ------------------------------- When was the issue noticed? Before use/during setup Please provide a detailed description of the event The received the sharps containers and no lids PHYSICIAN LAST NAME: Manthey PHYSICIAN FIRST NAME: cheryl PHYSICIAN CLINIC: FROEDTERT INTEGRATED SERVICE CENTER PHYSICIAN COUNTRY CODE: USA PHYSICIAN ADDRESS1: N86 W12999 Nightingale Way PHYSICIAN CITY: Menomonee Falls PHYSICIAN STATE CODE: WI PHYSICIAN POSTAL CODE: 53051 For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: bradln1 SALESREP EMP ID: 280982 SALESREP NAME: Bradley, Nicholas PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Bradley, Nicholas ADDITIONAL COMMENTS ABOUT THE EVENT: They received the sharps containers and now they need the lids free of charge As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2018-01-26 | 0702329973-10 | 8881112599 | 17K18763X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:INCORRECT ASSEMBLY | Note creator: Daniel Callahan Note Date: 02/20/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. A review of the device history record was performed with no anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112599 was 455 cases (800 parts per case). A total of five (5) samples were received with missing assemblies. Four (4) were missing plungers and rubber tips, and one (1) was missing the rubber tip. Samples were not contained in original packaging when received, but were shown via pictures from customer that product was inside original packaging with no sign of tampering/use. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The most probable root cause is tech operators not trained on how to handle RAM assembly equipment during production of Spraymation test samples per LS20001213. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time. The following corrective actions were implemented: 1. Training LS20001213 added to all Tech Operators on how to prepare samples for Spraymation testing. 2. Update Line 1 Checklist DC0006949 through Change Order # EC100737. Final Inspection sensors (plunger and tip presence) will be added to be inspected at beginning and end of each shift. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702329973 |  | Y | United States | Closed | Customer reports of incomplete assembled syringes and Large amount of syringes were missing plunger assembly | AQUABILITI | 0012728353 | 2018-01-25 | 2018-02-06 |  | 2018-01-24 | During use | 2018-05-12 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: QA Engineer - Mitchell Eichas ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: Large amount of syringes missing plunger assembly. Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: Still at Aquabiliti CONTACT NAME AND PHONE: Mitchelle Eichas REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event LOT# 17K18763X - unassembled - incomplete assembled syringes. Was the customer notified the product should be returned? No Please explain Waiting for investigation SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires written close-out letter Samples being sent and will be forwarded to Decontam Lab As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Veronica Dasari | Daniel Callahan | 2018-02-16 | 2018-03-01 | Note creator: Daniel Callahan Note Date: 02/16/18 Note Description: Addition of Analysis A total of five (5) samples were received with missing assemblies. Four (4) were missing plungers and rubber tips, and one (1) was missing the rubber tip. Samples were not contained in original packaging when received, but were shown via pictures from customer that product was inside original packaging with no sign of tampering/use. Samples were visually inspected per Quality Inspection Standard: Syringe Assembly and reported defect was confirmed for missing components. | Complete | #N/A
2018-01-26 | 0702330073-10 | 8881112599 | 17L16663X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12ML LLK SMOOTH TIP BULK NS | Unknown |  | 8881112599 |  |  |  | RFR | ASSEMBLY ISSUE | AFC | KC:MISSING COMPONENTS | Note creator: Daniel Callahan Note Date: 03/02/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. A total of seven (7) samples were received with missing assemblies. Five (5) had inverse rubbert tips not attached to plunger rod, and two (2) were missing the rubber tip. Samples were not contained in original packaging when received. Damage observed on finger flange of all five (5) assemblies with inverse rubber tips. A review of the device history record was performed for Lot# 17L16663X, with no applicable anomalies found. Lot size for part 8881112599 was 624 cases (800 parts per case), 499,200 total parts, on Line 1 were produced on 11/17/17. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). The most probable root cause is last check sensors for missing rubber tip and inverse rubber tip not functioning as intended. The following corrective actions implemented: 1. Update Line 1 Checklist DC0006949 through Change Order # EC100737. Final Inspection sensors (plunger and tip presence) will be added to be inspected at beginning and end of each Lot. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702330073 |  | Y | United States | Closed | Customer reports that rubber plunger tip is missing from 100 syringes | AQUABILITI | 0012728353 | 2018-01-25 | 2018-02-06 |  | 2018-01-24 | During use | 2018-05-12 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Quality Engineer - Mitchell Eichas ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112599, Description:8881112599 12ML LLK SMOOTH TIP BULK NS IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 8881112599 12ML LLK SMOOTH TIP BULK NS Model Number: 8881112599 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------------- Please provide a detailed description of the event LOT# 17L16663X - rubber plunger tip missing from 100 syringes. SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer requires written close-out letter Samples being sent As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Shanika Hodge | Daniel Callahan | 2018-03-01 | 2018-03-02 | Note creator: Daniel Callahan Note Date: 03/02/18 Note Description: Addition of Analysis A total of seven (7) samples were received with missing assemblies. Five (5) had inverse rubbert tips not attached to plunger rod, and two (2) were missing the rubber tip. Samples were not contained in original packaging when received. Damage observed on finger flange of all five (5) assemblies with inverse rubber tips. Could not determine if damage was incurred at customer or crystal lake facility. Samples were visually inspected per Quality Inspection Standard: Syringe Assembly and reported defect was confirmed for missing components. | Complete | #N/A
2018-02-09 | 0702359684-10 | 8881112083 | 17K15963X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112083 SYR 112 S-PLK SP X800 | MITG - Needles & Syringes | 003208 | 8881112083 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:MOLDING FLASH | Note creator: Daniel Callahan Note Date: 03/09/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. No physical samples have been returned for evaluation. The picture shown is the only one provided from customer. As shown in the image, there is a hair-like piece of material hanging from the left shoulder of the barrel. A review of the device history record was performed for Lot# 17K15963X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112083 was 240 cases (800 parts per case), on Line 7 were produced from 12/01/17 to 12/03/17. The method of root cause analysis that was implemented in this investigation was to conduct a 6M analysis that addressed: Man, Mater ial, Method, Measurement, Machine, and Mother Nature (Environment). Tool #58100 producing barrels with micro flash located on shoulder barrels. Friction between shoulder and tubing system creating angel hair (string flash). Maintenance Work Order CL-2775349 opened to address angel hair string flash on Tool #58100. This complaint will be used for tracking and trending purposes. If additional information is received warranting further analysis, the investigation may be resumed. | 0702359684 |  | Y | United States | Closed | The customer reports: The syringe had angel hair. | BUSSE HOSPITAL DISPOSABLES | 0012461165 | 2018-02-07 | 2018-02-15 |  | 2017-12-29 | During use | 2018-05-25 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Brooks, Janet PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Mary Sherman - Buyer ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112083, Description:8881112083 SYR 112 S-PLK SP X800 IMPLANTABLE: No WHO HAS POSSESSION: Other WHO HAS POSSESSION OTHER: STiil at Busse Hospital Disposables CONTACT NAME AND PHONE: Mary Sherman, Ph: 631-435-4711 REPLACED BY MDT PRODUCT: No, will be replaced in the future by a Medtronic product PRODUCT EVENT DETAILS: Event Details for 8881112083 SYR 112 S-PLK SP X800 Model Number: 8881112083 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------- Please provide a detailed description of the event LOT# 17K15963X. 5 out of every 50 syringes had angel hair. Customer rejected 35 cases. Was the customer notified the product should be returned? No Please explain Customer wants to return 35 cases - but need to wait for Plant to finish investigation SALESREP USER ID: brookj16 SALESREP EMP ID: 228025 SALESREP NAME: Brooks, Janet PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Brooks, Janet ADDITIONAL COMMENTS ABOUT THE EVENT: Customer thought she had sent this same issue but for item 8881112083 when she submitted for complaints 702237108,702237235 and 702237265. Customer wants RGA for 35 cases. Customer requires written investigation results. As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | Juan Diaz | 2018-05-14 | 2018-05-14 | Product Analysis was created by the Mansfield PMV Lab as part of the receipt process, so that Decontamination could be documented in GCH. This Product Analysis activity is being marked as “Complete” and the investigating site will create another Product Analysis activity upon receipt. | Complete | #N/A
2018-02-13 | 0702365211-10 | 1522SA | 17K08163 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 1522SA 2 QRT RED CONTAINER PHLEBOTOMY | MITG - Sharps Disposal | 002927 | 1522SA |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 1522SA evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for sample evaluation. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 17K08163 (12/04/17 to 12/20/17) and no internal reject found during production for the same issue. Trending Period: Feb 13, 2017 to March 14, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and two other complaints reported for the container size/lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging and ship to end user. The case quantity for 1522SA is 60 eaches/case. As per compliant quantity report is 1 each. As per compliant attachment “mPXR-498192-2018-02-13.pdf” the product was repacked and distributed by third party. The case may be broken and repackaged outside of Cardinal Health for further distribution. This shows that the Cardinal Health case was broken into smaller quantity at the external distribution center for further distribution. This was resulted in the failure of placing proper quantity of lids by external distributor prior to ship to end user. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used. | 0702365211 |  | Y | United States | Closed | The customer reported lids did not come with containers |  |  | 2018-02-13 | 2018-02-16 |  | 2018-02-13 | During use | 2018-03-17 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Platenyk, Kyle PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Distributor ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:1522SA, Serial/Lot No:17K08163, Description:1522SA 2 QRT RED CONTAINER PHLEBOTOMY IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 1522SA 2 QRT RED CONTAINER PHLEBOTOMY Model Number: 1522SA Serial or Lot Number: 17K08163 ------------------------------------------------------------------------ ------------------------------------ Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Yes Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event lids did not come with containers Was the customer notified the product should be returned? No Please explain will be sending customer replacement lids FACILITY NAME: Humana Specialty Pharmacy FACILITY COUNTRY CODE: USA FACILITY ADDRESS1: 9843 Windisch Road FACILITY CITY: West Chester FACILITY STATE CODE: OH FACILITY POSTAL CODE: 45069 SALESREP USER ID: platek1 SALESREP EMP ID: 326455 SALESREP NAME: Platenyk, Kyle PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Platenyk, Kyle As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-13 | 0702365441-10 | 8881512852 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512852 12ML S-R SYR ONLY O/P | MITG - Needles & Syringes | 003208 | 8881512852 |  |  |  | RFR | NO INFORMATION | AFC | KC:EMPTY SECONDARY PACKAGE | Note creator: Daniel Callahan Note Date: 3/09/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. No sample was returned and no picture available. The customer described the container as sealed with no parts inside. A review of the device history record was unable to be performed. Typical lot size for part 8881512852 is 360 cases (172,800 pieces) and production line 2/3 is used. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment that evaluated potential causes. The most probable root cause is mishandling of containers during clearing of auto packer jams/picker faults due to human error. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition presents a low risk level. This complaint will be used for tracking and trending purposes. | 0702365441 |  | Y | United States | Closed | The customer reported: Opened the case to pull out a box of syringes, one of the boxes is completely empty. It is still sealed | ANIMAL HEALTH CARE SPECIALIST | 0012933127 | 2018-02-13 | 2018-02-13 |  | 2017-12-13 | Unknown | 2018-03-14 | N | N | N | N | N | N |  | N |  |  | Null |  |  | Veronica Dasari | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-15 | 0702370262-10 | 31378089 | 00174720 | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 31378089 19 GAL SHARPS W/GASKET REDX5 | MITG - Sharps Disposal | 002927 | 31378089 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 31378089 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was performed and met acceptance criteria. DHR Review: Device History Records data reviewed for 174720 (01/22/17 - 01/24/17) and no internal reject found during production. Trending Period: Feb 06, 2017 to March 07, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and no other complaints reported for the product family/lid type. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 31378089 is 5 eaches/case. DHR reviewed and it passed all requirement as per specification. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702370262 |  | Y | United States | Closed | The customer reported they received case with missing lids | WEST MERCURY MINOR | 0012670550 | 2018-02-06 | 2018-02-16 |  | 2018-02-06 | Servicing | 2018-03-10 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Charlesworth, Victoria PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Office Manager at Hospital ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:31378089, Serial/Lot No:00174720, Description:31378089 19 GAL SHARPS W/GASKET REDX5 IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 31378089 19 GAL SHARPS W/GASKET REDX5 Model Number: 31378089 Serial or Lot Number: 00174720 ------------------------------------------------------------------------ -------------------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Received case with missing lids SALESREP USER ID: charlv2 SALESREP EMP ID: 278791 SALESREP NAME: Charlesworth, Victoria PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Charlesworth, Victoria As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-15 | 0702370264-10 | 8934PG2 | UNKNWON | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 12GA PGII SLIDE CHEMO - YELLOW | MITG - Sharps Disposal | 002927 | 8934PG2 |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8934PG2 evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Feb 13, 2017 to March 14, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and seven other complaints reported for the container size/different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to packaging. The case quantity for 8934PG2 is 10 eaches/case. As per the complaint attachment 1 case was missing lids but it was unclear how many lids are missing from one case. As the information provided is very limited and there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. | 0702370264 |  | Y | United States | Closed | The customer reported they received 2 cases of product and one case was missing lids. | RIVERSIDE CANCER INFUSION CENTER | 0012876927 | 2018-02-13 | 2018-02-16 |  | 2018-02-13 | Servicing | 2018-03-17 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Charlesworth, Victoria PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Office Manager at Hospital ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8934PG2, Description:12GA PGII SLIDE CHEMO - YELLOW IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 12GA PGII SLIDE CHEMO - YELLOW Model Number: 8934PG2 Serial or Lot Number: null ------------------------------------------------------------------------ -------------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? No Was the IFU for this product received and followed by the user? Yes When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Received 2 cases of product and one case was missing lids. SALESREP USER ID: charlv2 SALESREP EMP ID: 278791 SALESREP NAME: Charlesworth, Victoria PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Charlesworth, Victoria As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-16 | 0702373602-10 | 8989CRH | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 18 GAL CHEMO RED HINGE LID | MITG - Sharps Disposal | 002927 | 8989CRH |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8989CRH evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Feb 16, 2017 to March 16, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and fourteen other complaints reported for the product family with different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8989CRH is 5 eaches/case. As per the complaint attachment only 1 each is missing lid which show there was breakdown of product case in the supply chain. It is unknown from the complaint data if the product is repackaged at third party or distributor. As the information provided is very limited and there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. As per complaint there was no patient involvement which shows no harm or patient safety with missing lid. | 0702373602 |  | Y | Canada | Closed | The customer reported missing lid from sharps container, no patient involvement. |  |  | 2018-02-16 | 2018-02-21 |  | 2018-02-04 | Prior to Use | 2018-03-24 | N | N | N | N | N | N |  | N |  |  | Null |  | No Product found FIRST NOTIFIED NAME: Russell-Mondor, Joanne Event Context: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: materials management ASSOCIATED WITH PATIENT: No PRODUCT EVENT DETAILS OTHER: missing lid from sharps container, no patient involvement, no healthcare or doctor attached Below are the details for Product - Description:8989CRH IMPLANTABLE: No REPLACED BY MDT PRODUCT: Unknown, but will follow up for more information PRODUCT EVENT DETAILS: Event Details for 8989CRH Model Number: Not specified Serial or Lot Number: null ------------------------------------------------------------------------ --------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Unknown, but will follow-up for more information Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Unknown Unknown, but will follow-up for more information Was there any harm to anyone other than the patient? Unknown Unknown, but will follow-up for more information Please provide a detailed description of the event missing sharps container lid, no of these questions above pertain to this case Was the customer notified the product should be returned? Yes FACILITY NAME: St. Peter's Hospital FACILITY COUNTRY CODE: CAN FACILITY ADDRESS1: 20 Heritage Drive FACILITY CITY: Melville FACILITY STATE CODE: SK SALESREP USER ID: russej2 SALESREP EMP ID: 331012 SALESREP NAME: Russell-Mondor, Joanne PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Russell-Mondor, Joanne As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-16 | 0702373878-10 | 8939CRS | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 18 GAL CHEMO RED SLIDE LID | MITG - Sharps Disposal | 002927 | 8939CRS |  |  |  | RFR | PACKAGING | AFC | KC:MISSING LIDS | The Product 8939CRS evaluated for the missing lids. The reported condition cannot be observed as no photo or physical sample available for product analysis. DHR Review was not performed as lot number was not provided. DHR Review: A Device History Record was not reviewed as no lot number provided. Trending Period: Feb 16, 2017 to March 16, 2018 Trending Results: A review of internal inspection rejects for this particular issue was performed as well as a review of customer co mplaints covering the last 12 months. Results showed no internal inspection reject and fourteen other complaints reported for the product family with different lid style. Root Cause: A probable root cause has been determined to be human error and failure to verify the required amount of lids prior to ship to end user. The case quantity for 8939CRS is 5 eaches/case. As per the complaint attachment only 1 each is missing lid which show there was breakdown of product case in the supply chain. It is unknown from the complaint data if the product is repackaged at third party or distributor. As the information provided is very limited and there is no lot number provided so DHR cannot be reviewed. Formal Investigation: Based on the existing controls, the internal reject and the complaint history review, no formal investigation is required at this time. This will be used for tracking and trending purposes. Conclusion: This issue has been determined to be an event with minimal risk as missing lids product should not be used and must be replaced. As per complaint there was no patient involvement which shows no harm or patient safety with missing lid. | 0702373878 |  | Y | Canada | Closed | The customer reported missing lids for this product, missing lid form filled out by customer. | Victoria General Hospital | 0008036175 | 2018-02-16 | 2018-02-21 |  |  | During use | 2018-03-24 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Mavi, Sinem EVENT DATE: Unknown, but will follow up for more information PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - Sharps Safety ----------------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Healthcare Professional ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8939CRS, Description:18 GAL CHEMO RED SLIDE LID IMPLANTABLE: No REPLACED BY MDT PRODUCT: Yes PRODUCT EVENT DETAILS: Event Details for 18 GAL CHEMO RED SLIDE LID Model Number: 8939CRS Serial or Lot Number: null ------------------------------------------------------------------------ ---------------------- Was this product repackaged and sold by a third party (i.e. repackaged by distributor)? Unknown Unknown, but will follow-up for more information Was the IFU for this product received and followed by the user? Unknown Unknown, but will follow-up for more information When was the issue discovered? Prior to use Was there any harm to anyone other than the patient? No Please provide a detailed description of the event Missing lids for this product, missing lid form filled out by customer. Lids will be provided. Was the customer notified the product should be returned? No Please explain The product does not need to be returned. Missing lid form filled out. PHYSICIAN LAST NAME: Aylesworth PHYSICIAN FIRST NAME: Robert PHYSICIAN CLINIC: Victoria General Hospital PHYSICIAN COUNTRY CODE: CAN PHYSICIAN CITY: Victoria PHYSICIAN STATE CODE: BC For Physician : PHYSICIAN COMMUNICATION REQUESTED: No SALESREP USER ID: prokok1 SALESREP EMP ID: 252524 SALESREP NAME: Prokopy, Kimberly PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Prokopy, Kimberly As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Priyanka Ghule | 0 | 0 | 0 | 0 | Complete | #N/A
2018-02-22 | 0702383934-10 | 8881512878 | UNKNOWN | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881512878 12CC S-LK SYR ONLY O/P X480 | MITG - Needles & Syringes | 003208 | 8881512878 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:BROKEN LUER TIP | Reference attached investigation in GCH for a full concise summary. One sample was received. Sample was outside of its sterile hard-pack casing, indicating customer had broken heat seal. Luer tip was broken off the base of the sample and not present for evaluation. Lot number could not be determined. Typical lot size is 172,800 parts. Due to the part being an SR tip within sterile casing, it was produced on Line 23. A lot will not be released unless all acceptance criteria per established sampleing levels are within acceptable limits. he method of RCA was use of a 6 Why assessment. Most probable root cause was damage during aidlin transportation. Based on the infoprmation available a formal investiagtion is not deemed necessary at this time. The complaint will be used for tracking/trending. | 0702383934 |  | Y | United States | Closed | The customer reported: Hub was missing from this syringe, even though it was sealed in its sterile casing. | NORTH CASCADE VET HOSPITAL | 0012942049 | 2018-02-20 | 2018-02-22 |  | 2018-02-13 | Unknown | 2018-04-06 | N | N | N | N | N | N |  | N |  |  | Null |  |  | Sunitha Krishnanpillai | Daniel Callahan | 2018-03-27 | 2018-04-04 | 0 | Complete | #N/A
2018-02-23 | 0702385332-10 | 8881112059 | 15D00463X | COV LP/LLC, Crystal Lake, Mfg | PATIENT MONITORING & RECOVERY-PATIENT CARE |  | 8881112059 12ML SYRINGE REGULAR | MITG - Needles & Syringes | 003208 | 8881112059 |  |  |  | RFR | NO PATIENT INVOLVEMENT | AFC | KC:BROKEN LUER TIP | Note creator: Daniel Callahan Note Date: 03/21/18 Note Description: Addition of Investigation Summary Reference attached investigation for a full concise summary of all investigations performed. No physical sample has been returned for evaluation. One image was received for review. The syringe had been sealed in new packaging by Cook Medical and missing luer tip. A review of the device history record was performed for Lot# 15D00463X, with no applicable anomalies found. A lot will not be released unless all acceptance criteria inspections per established sampling levels are within acceptable limits. Lot size for part 8881112059 was 90,000 parts (500 parts per case), produced on Line 7 on 04/15/15. The method of root cause analysis that was implemented in this investigation was to conduct a Six M assessment. Probable Root Cause: Aidlin transport system damaged barrel during movement into kicker-wheel. Based on the information available, trend analysis, and the investigation findings, a formal investigation is not deemed necessary a t this time as the reported condition presents a low risk level. This complaint will be used for tracking and trending purposes. | 0702385332 |  | Y | United States | Closed | The customer reports: Before opening the package, the user found the tip of the syringe in the set was broken and missing.This issue was noticed prior to patient contact. | Blocked- Cook Medical | 0001276800 | 2018-02-22 | 2018-02-23 |  | 2018-02-20 | Prior to Use | 2018-03-24 | N | N | N | N | N | N |  | N |  |  | Null |  | FIRST NOTIFIED NAME: Massey, Joanne PATIENT EVENT DETAILS: Was a Medtronic person present at the event? No Patient Event Details for MITG Patient Care - OEM ------------------------------------------------- None specified If applicable, please provide additional details about what happened to the patient relevant to this event ------------------------------------------------------------------------ ---------------------------------- Not specified PATIENT GENDER : %MdtMPXR_Complaint.docs:ComplaintData/report/patient/Pa tient.genderXlate/*body% Complaint Source: Other COMPLAINT SOURCE OTHER: Cook Medical ASSOCIATED WITH PATIENT: No Below are the details for Product - Model No:8881112059, Description:8881112059 12ML SYRINGE REGULAR IMPLANTABLE: No REPLACED BY MDT PRODUCT: No, will not be replaced in the future PRODUCT EVENT DETAILS: Event Details for 8881112059 12ML SYRINGE REGULAR Model Number: 8881112059 Serial or Lot Number: null ------------------------------------------------------------------------ ------------------------------ Please provide a detailed description of the event The product was returned to Cook Medical unused and unopened in its original packaging and It was found that the top of the syringe tip was missing/ broken off the syringe. 1 unit Lot # 15D00463X Was the customer notified the product should be returned? No Please explain Customer only had 1 unit that was defective pictures of the defect was sent SALESREP USER ID: massej1 SALESREP EMP ID: 228168 SALESREP NAME: Massey, Joanne PRIMARY CONTACT: Field Contact listed above COMPLAINT OWNER NAME: Massey, Joanne As Reported Event - interfaceText: Please confirm that all products or components which are part of the system related to this complaint are included above. ------------------------------------------------------------------------ ------------------------------------------------- Yes | Sunitha Krishnanpillai | 0 | 0 | 0 | 0 | Complete | #N/A